Bluepoint Medical repair codes - pulse oximeters

Applies to:

0012160

0012161

0012162

0012163

0012164

SMARTsat

0012165

0012166

0012167

| Category | Description |
| --- | --- |
| Display - dark | Aged OLED got dark |
| Display - defective | Defective display e.g. TFT backlight defective, mechanical defect due to drop |
| Fluid damage | Damage to PCBA of battery contacts due to ingress of fluid |
| Housing damage | Mechanical damages to the housing |
| Keypad damage | Mechanical damages to the keypad |
| Main PCBA defective | Issue was solved by exchanging the main PCB |
| Power PCBA defective | Issue was solved by exchanging the power PCB |
| SpO2 flex socket | SpO2 flex socket connecting to the main PCB is defective typically resulting in “sensor defective” error message |
| SMARTsat defective | SMARTsat module defective. Detailed analysis of SMARTsat defects performed in a separate analysis |
| ChipOx defective | ChipOx module defective |
| SpO2 Sensor LED defective | SpO2 Sensor defective message. Analysis of error channel shows the LED (red or ired) is defective. Mostly due to cable break issues. |
| SpO2 Sensor detector defective | SpO2 Sensor defective message. Analysis of error channel shows the detector is defective. Mostly due to cable-break issues. |
| Wrong SpO2 sensor | SpO2 Sensor defective message. Analysis of error channel shows the wrong sensor is used. The sensor ID is not compatible with the monitor |
| Device SW Update | SW Update of the monitor solved the issue |
| SMARTsat SW Update | SW Update of the SMARTsat SpO2 Module solved the issue |
| Other | e.g. ChipOx came loose from main pcb |
| Not reproducible | The defect is not reproducible |