Bluepoint Medical repair codes – capnographs, IRMA, ISA

Applies to:

4410500

4410501

4410520

4410521

4410530

4410540

4410550

4410560

4420505

4420506

4420507

4420508

| Category | Description |
| --- | --- |
| Display - defective | Defective display e.g. dark OLED, mechanical defect due to drop, pixel error |
| Fluid damage | Damage to PCBA of battery contacts due to ingress of fluid |
| Housing damage | Mechanical damages to the housing |
| Keypad damages | Mechanical damages to the keypad |
| Main PCBA defective | Issue was solved by exchanging the main PCB |
| Power PCBA defective | Issue was solved by exchanging the power PCB |
| SpO2 wire socket | SpO2 wire socket connecting to the main PCB is defective typically resulting in “sensor defective” error message |
| ChipOx defective | ChipOx module defective |
| SpO2 Sensor defective | SpO2 Sensor defective (defective photodiode, LED, wrong sensor ID etc.) |
| Device SW Update | SW Update of the monitor solved the issue |
| Power supply defective | Issue was solved by exchanging the power supply |
| IRMA defective | Issue was solved by exchanging the IRMA (detailed analysis based on Masimo failure report in case available) |
| ISA defective | Issue was solved by exchanging the ISA (detailed analysis based on Masimo failure report in case available) |
| Li-Ion battery defective | Issue was solved by exchanging the Li-Ion battery |
| Other | e.g. ChipOx came loose from main pcb |
| Not reproducible | The defect is not reproducible |
| Functional maintenance check | No defect, annual function maintenance test upon customer request |