Bluepoint Medical repair codes – SMARTsat sensors

0014651

0014752

0014753

0014754

0014760

0014850

0014851

0014895

0014896

0015010

0015011

0015012

0015013

| Category | Description |
| --- | --- |
| SpO2 Sensor LED defective | SpO2 Sensor defective message. Analysis of error channel shows the LED (red or ired) is defective. Mostly due to cable break issues. |
| SpO2 Sensor detector defective | SpO2 Sensor defective message. Analysis of error channel shows the detector is defective. Mostly due to cable-break issues. |
| Wrong SpO2 sensor | SpO2 Sensor defective message. Analysis of error channel shows the wrong sensor is used. The sensor ID is not compatible with the monitor |
| Connector defective | Mechanical defect at the connector |
| Housing defective | Mechanical defect at the housing |
| Cable defective | Mechanical defect at the cable |