QC 21 Non Conformance Report
	Date
	11/11/2021

	Viamed Issue id
unique identifier 
	241894

	BSI Ref (if applicable) 
unique identifier 
	2128198-202111-N3 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	[image: image1.png]Assessment Report.
Finding ¥ Certificate
2 2128198-202111-N3 D 78787
Certificate 150 134852016 & EN 1SO
Standard 13485 e 029
Tocation
0009370214-000
“Assessment
e 3000498
Category Minor
Feedback Processes, Complaint Management, Vigiance and
Area/process: | Fe=Cpeck P
Statement of | Reporting to regulatory authorites was not fully effective as procedure
non- s not seen to be updated according appiicable reguiatory.
conformance: _| requirements.
Reporting to regulatory authorities
I applicable reguiatory requirements require notifcation of complaints.
that meet speciied reporting citeia of adverse events or issuance of
E=o ‘advisory natices, the organization shall document procedures for
requirements | oro,iing notification to the approprate regulatory authorites.
Records of reporting to reguiatory authorites shall be maintained (see
425).

o ‘Compaints Vigilance and NoGfcation Viamed Ltd VOP19, 30. 08. 2019
EED Reporting time frame and reporting form has not been updated.
Cause
Correction/containment
Corrective action
Finding o Certificate

1 2128198-202111-N4 D 78787
Certificate 150 134852016 & EN 150
Standard 13485 G &3
Location
0009370214-000
“Assessment.
= 3000498
Category Minor
Improvement - Preventive and Corrective Actions incuding Process
Area/process: | change Management





 

	Investigation By:
Person responsible 
	Derek Lamb

	Root Cause Analysis Investigation Issue id
(if applicable)
	Our document VOP 19 had not been updated with the newest Reporting time frame and reporting form, for the transition between MDD and MDR. As all of our own products had been discontinued in 2019, we had not realised the documentation needed to be updated for the new MDR.


	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action 
Issue ID (if applicable):
Relevant Immediate Corrective action (if applicable)
	We will update the VOP 19 to reflect the correct and current reporting time frame for the MDR. 
We will update the VOP 19 to reflect the correct reporting method / form for the MDR.


	Time Scale for Immediate Corrective Action
Time for completion of all identified actions  
	7 Days from now, 18th November 2021

	Corrective Action:
Relevant and Proportionate Corrective Action
	We will review any documents that have reference to the MDR and make sure they are up to date. 
To prevent this happening again we will add a rolling task to review the MDR and standards in relation to our discontinued products, to make sure this review is done regularly. So we do not miss any future updates. 


	Time Scale for Corrective Action
Time for completion of all identified actions 
	30 days from now, 11th December 2021


	Risk Review 
Confirm action does not have negative affects and/or risks to any other part of the ISO system 
	The update to the VOP 19 does not pose any risk to the system or processes, as it is an improvement to the system.
The process we have added does not pose any risk or negative affects, to the system or to other processes. 
See Issue 241894 for full Risk Review
Risk Assessment For Updating Document ID31040 


	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


