QC 21 Non Conformance Report
	Date
	11/11/2021

	Issue id
unique identifier 
	241920

	BSI Ref (if applicable) 
unique identifier 
	2128198-202111-N8 

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 

	 [image: image1.png]Finding Certificate
o e 2128198-202111-N8. MD 78787
Certificate 150 13485:2016 R EN 1SO
eeadid TS Clause 414
Location
s 0009370214-000
“Assessment
o 3000498
Category Minor
Area/process: _| Control of Documents and Records
Statement of | Control of change was not fully effective as relevant procedure and
non- ‘system (Intrastats) didn' include change impact assessment,
conformance: _ | determining affected areas/documents and if change is significant.
“The organization shall manage these quality management system
processes in accordance with the requirements of this International
) ‘Standard and applicable regulatory requirements. Changes to be made.
= to these processes shall be:
requirements | ;) evaluated for their impact on the quaity management system;
b) evaluated for their impact on the medical devices produced under this.
quality management system;

making excellence a h

Assessment Repor

0) ontrolled n accordance with the requirements of Gis International
Standard and applicable regulatory requirements.

VOPOI issue 1, 30. 09. 2019
Intrastats change control module.
‘sample: Change number 127253, 3. 09. 2018, 27436, VOPL

/containment






	Investigation By:
Person responsible 
	Derek Lamb


	Root Cause Analysis Investigation Issue id
(if applicable)
	We were unable to provide adequate proof that we were assessing the impact of changes, to our System documents, processes or to our existing product documentation, that relate to products that are still within their life span. 


	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action 
Issue ID (if applicable):
Relevant Immediate Corrective action (if applicable) 
	We will update our top level document VOP 01 Documentation and Records, to include the statement - 
‘Any change to system documentation or process(s), if they are linked to either a VOP, ISO Standard, Sub Process, or Controlled Product Technical File, will have a Risk of Change or Impact Assessment created. 
This assessment will then be uploaded to the controlled documents in the Document Index and linked to Top Level Document that is to be updated. ‘


	Time Scale for Immediate Corrective Action
Time for completion of all identified actions
	10 Days, 23rd November 2021

	Corrective Action:
Relevant and Proportionate Corrective Action
	We will update the system to ensure that any changes made to the System or Process documentation, will force a Risk / Impact Assessment of the change to be carried out, in the related areas of the System, process or controlled product Technical File, before the change can be enacted. 
This corrective action will prevent us from making changes, to documents, system processes or controlled product technical files, without carrying out an assessment of risk or negative affect. 


	Time Scale for Corrective Action
Time for completion of all identified actions 
	30 days from today, 14th December 2021

	Risk Review 
Confirm action does not have negative affects and/or risks to any other part of the ISO system 
	There is no risk in the update of the system, as this is an extra preventative measure. 
This action does not have a negative affect and/or risks to this or any other part of the ISO system
See Issue 241920 file
Risk Assessment For Updating Document ID30999 VOP 01
 

	Follow-up future issue id
(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


