en 	Instructions For Use for reusable Temperature probes

[Type list …]

Intended Use

The reusable temperature probes are intended to be used for monitoring temperature. The temperature probes are designed for use with GE Healthcare monitoring systems (e.g. CARESCAPE B450, B650 and B850 monitors) and other monitors compatible with YSI 400 series temperature probes.

The temperature probes are indicated for use by qualified medical personnel only.

Warning

- Improper use of probes can lead to damage of the internal wiring and to the loss of electrical insulation as well as incorrect temperature readings. In case of probe damage, the probe must not be used.

- Do not apply the skin probe on areas with open wounds or lesions.

- This temperature probe is intended for use in combination with compatible monitors only. Usage in combination with other devices can lead to malfunction.

- Strong sources of electromagnetic fields, e.g. electro-surgical devices, can adversely affect the functioning of the equipment. If possible, the probe should be completely removed from the patient before a surgical device or other HF source is activated! Check the protective earthing if you intend to use HF devices and temperature measuring equipment in parallel. Minimize the potential risk by selecting a temperature measurement point far away from the HF current path.

- MR unsafe. We do not recommend the use of the probes with magnetic resonance imaging (MRI) or computed tomography (CT) devices.

- Stop using the probe immediately if it becomes visibly damaged, including pits or cracks in the surface of the probe or connector, or exposed wires.

- For further information and warnings please read the usage instructions accompanying the temperature-measuring device.

Instructions for Use

- Temperatures can be measured e.g. in the esophagus, on the skin surface or rectally. Please use water-based lubricants only. Maximum continuous contact duration for rectal and esophageal application is 24 hours, and for skin contact application less than 30 days.

- Skin: Place the skin probe disk in direct contact with intact skin using medical adhesive tape. In special circumstances, the extremities such as fingers or toes may also be used. Make sure that movement does not interrupt the skin contact.

- Esophagus: The probe should be placed in the lower third of the esophagus.

- Rectally: The probe is in correct position when it has passed the sphincter.

- Thoroughly clean immediately after use to prevent soil from drying on the probe – refer to reprocessing instructions below.

Specifications

Operating mode:					Direct mode

Overall length:					280cm (nominal)

Temperature range:				Storage/Transport: 	-30°C…+70°C

Operation: 		+0°C…+50°C

Humidity range:					Storage/Transport:	10… 95% r. h. non-condensing

Operation:		10…100% r. h.

Ambient pressure:				Storage/Transport:	800…1060hPa

Operation:		800…1060hPa

Accuracy:					± 0.1°C (range 25°C…45°C)

Transient response time (2 K transient):		Large types:		< 50s

Medium types:		< 30s

Small types:		< 25s

Reprocessing instructions

| Initial treatment: | Before cleaning or disinfecting the probe, it has to be detached from the monitor. Special handling during transport is not necessary. For removal of visible soil see under 'Manual cleaning' or ‘Machine cleaning’. Damage to the outer cable jacket of the probe must be avoided (for example by sharp items during transportation). |
| --- | --- |
| Preparation before cleaning: | The probe does not contain any parts to be disassembled before cleaning. Check the probe for damage before cleaning (visual inspection). If damage is found, take the probe out of operation and dispose of it according to local regulations for medical waste. |
| Cleaning: Manual cleaning: Ultrasound cleaning: Machine cleaning: General: | To remove visible soil, clean the probe by wiping it with a cleaning agent and a soft cloth or by placing it in an immersion bath with cleaning agent (see below). Then rinse the probe sufficiently (softened or deionized water recommended) and let it dry in ambient air for at least 30 minutes. Follow the manufacturer's instructions when using the cleaning agent including dilution from concentrate. Avoid high mechanical force to the probe during cleaning process. Cleaning agents: Prolystica®: 15 min exposure time; max. 55°C Do not use ultrasound-cleaning methods on the temperature probes – this may damage the probe components. The probe can be cleaned by washing in a machine (cleaning and disinfection device) at up to 95 °C (cleaning agent e.g. neodisher® Mediclean Forte). Follow the manufacturer's instructions when using the cleaning agent and the automated washing equipment instructions for maximum load size. The recommended settings are: - Step 1 cleaning 0.5% cleaner 30°C for 1 min - Step 2 cleaning 0.5% cleaner 55°C for 12 min - Step 3 rinsing with demineralized water for 1 min Inspect probes after cleaning to ensure all visible soil is removed. If the device is determined not to be visually clean, repeat the cleaning steps or safely dispose of the device. |
| Disinfection: Chemical: Thermal: | High level disinfection (skin contact probes only): Disinfection by placing it in an immersion bath with disinfecting agent (e.g. Cidex® OPA) at 20°C for at a minimum of 12 minutes. Then rinse the probe sufficiently according to the manufacturer’s directions for use (high-purity or distilled water recommended) and let it dry in ambient air for at least 30 minutes. In any case, refer to the IFU supplied by the disinfectant manufacturer. Avoid the use of disinfectants containing ketones. Thermal disinfection at 90°C with 5 min exposure time possible. |
| Water Quality: | Use water quality appropriate to the reprocessing step, including for dilution preparation of agents (distilled water recommended), rinsing after cleaning, and particularly after high level disinfection (high purity or distilled water). Automated cleaner and/or disinfector machines may have specific water requirements as well. See also: AAMI TIR34, Water for the reprocessing of medical devices. |
| Drying: | Drying after cleaning, disinfection or sterilization for at least 30 minutes at room temperature. For more information on the drying process, see Cleaning and Disinfection. Do not dry the probe with hot air above 90 °C. |
| Inspection and maintenance: | If damage to the probe (e.g. cracks in the cable jacket, deformation of plastic parts or connectors) can be detected, the probe must not be used and shall be replaced. If the regular function with the medical device is not given (e.g. caused by open circuit), the probe must also be replaced as well. |
| Packaging: | Steam sterilization has been validated using ASSURE Plus sterilization pouches REF: 83005. We recommend using this sterile barrier system for packaging of the temperature probes prior to sterilization. Packaging in an appropriate pouch is also recommended immediately after the cleaning or disinfection process. |
| Sterilization: | The temperature probe can be sterilized with hot steam procedure (autoclave) at 132°C/10 min with a drying time of at least 20 minutes (gravity process), or at 132°C/3 min with a drying time of at least 16 minutes (forced vacuum process). Follow the generally applicable instructions and regulations. Clean the temperature probe before sterilization and disinfect the probe according to applicable reprocessing rules. |
| Reuse life: | The temperature probes have been validated for the following reprocessing cycles: Manual cleaning/disinfection: up to 300 cycles Machine washing/disinfection: up to 300 cycles Sterilization: up to 100 cycles Stop using the probe immediately if it becomes visibly damaged, including pits or cracks in the surface of the probe or connector, or exposed wires. |
| Storage: | Storage protected from dust in original packaging or comparable packaging. Avoid continuous direct sun light. Further information under Specifications. |
| Transport: | Transport to the place of use protected from dust in original packaging or comparable packaging. |

Maintenance/Service

The temperature probes do not contain serviceable or replaceable components. Modification or replacement of components is prohibited. It is recommended to run an inspection of accuracy at intervals of two years.

Returning the product

Every product sent to us for inspection must include a signed confirmation that the product has been cleaned and disinfected beforehand.

Disposal

Before disposing the product, it may be necessary to clean and disinfect it. Local regulations may apply for disposal of used medical devices.

neodisher Mediclean Forte is a trademark of Chemische Fabrik Dr. Weigert GmbH & Co. KG.

Prolystica is a tradermark of STERIS Corporation.

Cidex OPA is a trademark of ADVANCED STERILIZATION PRODUCTS, a Johnson+Johnson company

Follow instructions for use. Symbol appears blue on device.

For prescription use only (US)

Product is free of latex

Product is free of phthalates

Do not dispose in domestic waste!

Consult Operating instructions

IP44		Degree of protection

Catalogue number

Serial number

Manufacturer

Manufacturing date

Atmospheric pressure limitation

Humidity limitation

Temperature limitation

CE-Conformité Européene mark, 0482-notified body

China RoHS conform

Do not throw away

MR unsafe

Protect from direct sunlight

1 Piece

Product is supplied non-sterile

Sales Representative Switzerland

bluepoint medical GmbH & Co. KG

An der Trave 15, 23923 Selmsdorf

Germany