Investigation into VM-2160 SMARTsat in Neonatal Applications

Customer:	Grange University Hospital, Wales. CID19789
Issue raised by:	Clare Lewis, Technical Educator, Neonatal Intensive Care Unit
Equipment:	p/n 0012165 VM-2160 SMARTsat
with 0015013 Disposable Neonatal Sensor
Viamed Issue:	256351

Email received 4th April 2022 from, Clare Lewis:
“We recently purchased 6 of the handheld VM-2160 pulse oximetry monitors from you for use in our delivery room care and purchased the neonatal disposable plaster wrap around probes. Unfortunately, we have experienced problems with their use in theatres and the delivery rooms, where the saturations and heart rates are not being picked up and we are just given a bad signal warning. As I'm sure you can imagine the first few minutes of delivery room care is the most important and we are starting at a disadvantage because are new monitors are not picking up. The cost of buying more saturation monitors is not something that as a department we can facilitate and before we purchased these from yourselves, I was told of the quality and accuracy of the item. Please could you advise a way forward.”

I spoke to Sally Pyrah-Barnes, Clinical Ward Sister and Neonatal Nurse Technician with over 30 years experience in this field to discuss their experiences and request more details.

Background clinical information: Pulse oximetry is used routinely the delivery room to monitor transition from intrauterine to extrauterine life. This transition requires a rapid adaptation of multiple organ systems. Separation from the placental circulation results in increased systemic vascular resistance, while initiation of ventilation lowers pulmonary vascular resistance.

The customer reports an inability to get saturation readings during the transition period. At the request of Bluepoint, we asked a number of questions to get further information on the issue.

---

Question: how is the VM-2160 SMARTsat being used?

Response: the department uses hand-held monitors immediately after delivery to gain vital information for decision-making that “has massive implications for resuscitation”.

The pulse oximetry sensor is placed on the neonate’s right hand to gain pre-ductal readings in line with Neonatal Life Support Guidelines in the UK.

The device is used for initial assessment only; once the patient is established, they will be moved to full patient monitors if further monitoring is required.

---

Question: Describe in detail exactly what the problem is, including details such as whether you can get any reading at all, whether it is difficult to align the sensor and maintain the reading etc.

Response: no reading at all on any patients for the first 10-12 minutes during the transition, it is sometimes up to 30 minutes before a reading is obtained. This is not an alignment issue; no amount of realigning allows a reading to be obtained.

---

Question: approximately what proportion of patients are problematic to gain a reading on?

Response: no reading is obtainable on any patient until at least 10-12 minutes old.

---

Question: is there anything specific about the size or condition of difficult to monitor patients?

Response: no, they have a diverse range of patients from 470 g up to 5 kg and are unable to gain readings even on the larger patients. Skin colour is also not a factor.

---

Question: have they had any success with other oximeters in this application?

Response: yes, they have no problem using Nellcor Oximax devices in the exact same application, stating that Nellcor Oximax devices that “work every time”. To support the conclusion that this is not a patient-dependent issue, they used the VM-2160 and a Nellcor Oximax device on identical twins and found the same conclusions.

---

Question: this customer has been supplied with sensors from 2 different batches: lot numbers 077086 and 098544. Have they tried sensors from both batches, and if so, any difference in usability?

Response: as yet, she is insure which batches the sensors have come from but will arrange for them both to be tested and report the results.

---

Progress of investigation:

Customer has purchased 6 boxes of 0015013, which we received in 2 different batches:
Barcode IDs: 1704751, 1704752, 1704748 - Received 03 Mar 2021 on PVM1582
Barcode IDs: 1895208, 1895201, 1895202, PVM2039 - Received 24 Nov 2021.

25th May 22 – After we advised the customer with the batch data, they initially tried 2 from each batch and neither picked up a reading.

15th June 22 - Customer has now tried 4 more, 2 from each batch and found that none of them pick up in the time immediately after delivery. Added to issue #25635 to SN to ask how we proceed, whether Bluepoint are coming up with a solution or whether we should credit them.

16th August 22 - Steve Nixon has discussed with Bluepoint engineers, who are looking at options to resolve this and may require additional clinical data, which Viamed will help to facilitate where possible. Could be due to low perfusion, could be resolved by OEM II and/or sensor design change. Or could be due to algorithm that needs updating specifically for newborns.

17th August 22 - Steve Nixon: The users are not using the equipment in line with the current IFU for either the oximeter or sensor type. However, from the feedback we will investigate the feasibility of monitoring being carried out on the hands (especially the right hand) of newborns during the immediate half hour after birth. During this transition period there are changes in neonatal physiology that makes existing monitoring challenging to overcome.