Specification Datasheet

Viamed

SpiroTrue A(PC)

Flow Sensor

CE

| ## Part numbers | 4310002 pack of 5 |
| --- | --- |
| ## Model | SpiroTrue A(PC) |
| Primary packaging | PE pouch (bag) with 5 sets |

#### Technical Data

| Measuring principle | Hot-wire anemometric principle |
| --- | --- |
| Reproducibility | ≤ 5% at flow range 0 to 120l/min |
| Housing material | PC (Polycarbonate), clear transparent |
| Connector | 4 pin gold plated |
| Weight | ≤ 11g |
| Components conform to | 2011/65/EC (RoHS) |
| Conforms to | EN ISO 10993-5 |
| Designed in accordance with | 93/42/EEC (Annex I and Annex II) |
| Classification in accordance with | Annex IX of 93/42/EEC: class lla |
| Compatible to | Dräger Spirolog REF 8403735, SpiroLife REF MK01900 |
| Packaging | Non-sterile |
| Storage temperature | # -20ºC to 50ºC |
| Humidity (R.H.) | 5% to 95% |
| Storage period | ≤ 3 years (unopened, away from direct sunlight) |
| Operating temperature | 10ºC to 40ºC |
| Relative humidity | 5 to 95%, non-condensing |
| Notes: Before using the sensor consult the appropriate ventilator user instructions for more detailed information. Applications Devices without RFID functionality e.g. Evita family, Savina family, Fabius family, Primus/Apollo, Julian, Cato, Cicero Flow module A for SpiroTrue A/APC Sensor life The sensor can be used as long as the calibration procedure can be carried out in the anaesthesia or ventilation equipment. Sensor disinfection The following equipment disinfectants may be used: Korsolex Basic/Extra/Plus; Gigasept FF/AF; Incidin liquid Disinfect the flow sensor by immersion Sensor cable available separately Cable length: 500mm Cable OD: 5mm Cable jacket: Medical grade PVC, RAL 7032 Plug over moulded: Medical grade PVC, RAL 7032 |  |