Specification Datasheet

Viamed

SpiroTrue D

OEM Disposable Single Patient Adult Flow Sensor

CE

| ## Part numbers | 4310063 OEM pack of 20, 2.2m tubing variant 4310064 OEM pack of 20, 3.2m tubing variant |
| --- | --- |
| ## Model | SpiroTrue D – OEM disposable single patient adult flow sensor & tubing set |
| Primary packaging | PE pouch (bag) with 20 sets includes flow sensor with double PVC (DEHP free) tubing connected |

#### Technical Data

| Sensor Technical Data |  |  |
| --- | --- | --- |
| Measuring principle |  | Differential pressure measurement |
| Reproducibility |  | ≤ 10% at flow range 0 to 100l/min |
| Designed in accordance with |  | 93/42/EEC (Annex I and Annex II) |
| Housing material |  | ABS (Acrylnitril-Butadien-Styrol) |
| Weight |  | ≤ 9.5g (without tubing) |
| Dead space |  | ≤ 10ml |
| Connectors |  | 15mm O.D./15mm O.D. in accordance with ISO 5356-1 |
| Sensor conforms to |  | EN ISO 10993-5 |
| Sensor and tubing conforms to |  | 2011/65/EC (RoHS) |
| Packaged |  | Non-sterile |
| Recommended storage conditions |  |  |
| Storage temperature |  | # -20ºC to 50ºC |
| Relative Humidity |  | 5% to 95% |
| Storage period |  | ≤ 3 years (unopened, away from direct sunlight) |
| Recommended operating conditions |  |  |
| Operating temperature |  | 10ºC to 40ºC |
| Relative humidity |  | ≤ 95%, non-condensing |
| Notes: Before using the sensor consult the appropriate ventilator user instructions for more detailed information. Sensor was tested on OxyLog 3000 SW1.21 Setup: 150 Dräger circuit REF 5702871-14; 2.0l Dräger test lung with restrictor 7mm: REF 2165694 |  |  |
|  | Hook and adapters are available separately 4310014 Adapter 15mm I.D. – 15mm I.D. Pack of 200 4330014 Adapter 15mm I.D. – 15mm I.D. Single unit 4310015 Hook. Pack of 200 4330014 Hook. Single unit |  |