MD300-W Sleeping Monitoring Pulse Oximeter
The MD300-W Wrist Oximeter is a small, wrist-worn device indicated for use in measuring, displaying and storing functional SpO2 and pulse rate. It may be used for spot-checking, data collection and recording of adult and paediatric patients in hospitals, pre-hospital, homecare, subacute and sleep study environments.

Features

· Compact and lightweight design.

· Application from adult to neonatal.

· Low power consumption, 20 hours continuous working with 1 x AAA Alkaline battery.

· Data transmission, storage and replay.

· 1 – 10 patient ID setting.

· Low voltage indicator.

· Easy to transmit data to computer.

Standard Accessories

· 2 x AAA Alkaline Batteries.

· 1 x SpO2 Soft Silicone Finger Sensor.

· 1 x Reusable Wrist Strap.

· Operating Manual

· USB Data Cable

· Data Management Software CD

Technical Specifications:

· SpO2 Range 0 – 100%

· Measuring range: 70 – 100%

· Resolution 1%

· Accuracy
80 – 100%
±2%

70 – 79% 
±3%



  0 – 70 %
Unspecified

· Pulse Rate Display range: 0 – 254 bpm

· Pulse Rate Measuring range: 30 – 235 bpm

· Resolution: 1 bpm
· Measuring Accuracy:
  30 – 100 ±2 bpm
101 – 235 ±2%

· Display type: LCD
· Parameters: SpO2, Pulse Rate, Signal Strength indicator
Alarm

· Low power indicator, probe off, finger out.

· Data transmission.

Storage

· Data records stores and replay 72 hours.

· Operating Temperature: 5 – 40°C

· Operating Relative Humidity ≤ 80%
· Storage Temperature: -20 - 70°C
· Storage Relative Humidity ≤ 93%
Battery

· Type: 1 x AAA Alkaline Battery

· Operating time approximately 12 hours continuously.

Dimensions and Weight

Dimensions: 60 mm (L) x 50 mm (W) x 20 mm (H) 

Weight: 26 g
