Positioning Made Simple
with Blue Diamond® BeanBags

For numerous years Blue Diamond® Bean Bags have been used worldwide in hospitals and clinics to provide improved surgical positioning.

The products have been used effectively in thousands of procedures. When properly utilized they will aid in optimal surgical site positioning.


BeanBags are designed with 2 distinct purposes. To aid in precise positioning during surgical procedures; and to help reduce the risk of decubituae and sheer by offering more evenly distributed body mass and reducing pressure points at bony prominences.


Our BeanBags are available in many configurations as well as two distinct variations, traditional and with integrated Blue Diamond® viscoelastic polymer gel layer.


BD-BB3040  a 30” x 40” bladder with a ‘U’ shaped cutout for shoulder




Available in traditional and gel versions


BD-BB4036
a 36” x 40” bladder with a ‘U’ shaped cutout for shoulder




Available in traditional and gel versions


BD-BB4646  a 46” x46” bladder with a ‘U’ shaped cutout for shoulder




Available in traditional and gel versions


BD-BB3076  a 30”x76” bladder full length BeanBag


BD-BB1020  a 10”x20” bladder for specialized positioning




Available in traditional and gel versions


BD-BB2020
a 20”x 20” bladder for specialized positioning




Available in traditional and gel versions

Traditional BeanBags

Our Traditional Bean Bag is comprised of an extremely durable polyurethane shell that is designed to be fluid and stain resistant. It is a soft and pliable film that will mold and contour to the patients’ body as suction is applied to the bladder.

The interior of the BeanBag is filled with thousands of varied sized beads that when vacuum is applied will come together and form a firm positioner that will mold around the patient.

The valve system on our BeanBag is truly unique. It is a 2 part system than can be removed and replaced if damaged. This does not apply if the red valve stem is torn from the urethane bladder. 
!!This valve stem should not be used as a carry handle or to hang the BeanBag from during storage. It can result in product failure!!
Blue Diamond® Gel BeanBags

Our Blue Diamond® Gel BeanBag is comprised of an extremely durable polyurethane shell that is designed to be fluid and stain resistant. It is a soft and pliable film that will mold and contour to the patients’ body as suction is applied to the bladder.

On one side of the Beanbag is an independent layer of our Blue Diamond® viscoelastic polymer gel. This will truly aid in the prevention of decubituae and sheer incidences during surgical procedures using our BeanBag.


The interior of the BeanBag is filled with thousands of varied sized beads that when vacuum is applied will come together and form a firm grasping positioner around the patient.


The valve system on our BeanBag is truly unique. It is a 2 part system than can be removed and replaced if damaged. This does not apply if the red valve stem is torn from the urethane bladder. 

!!This valve stem should not be used as a carry handle or to hang the BeanBag from during storage. It can result in product failure!!

Suggested Protocol:
1. Place BeanBag on the surgical table. Smooth out and evenly distribute beads throughout the bladder.

If using the Gel version:
Gel side must be facing up towards patient in order to be effect in the aid of decubituae and sheer prevention. BeanBag will function will Gel side facing out but will not provide the added benefit of the Gel layer.

2. The valve should be open so air can travel throughout the bladder.

 Using your current protocol, place the patient onto the surgical table on top of the BeanBag.

3. Again using your current positioning protocol,( we suggest at least 2 persons) mold the BeanBag around the patient while leaving the surgical site exposed so no part of the BeanBag will interfere with the surgery.
4. When optimal position is achieved suction should be applied to the valve system. The time to have the BeanBag become rigid will depend on the size of the bladder and suction power at your facility, usually the BeanBag molds very rapidly(within several minutes). We suggest evacuating as much air as possible to form a firm bladder that will hold and support the patient. 
*You cannot “over suction” as once all air is removed from the bladder it will not become anymore rigid.
5. Once optimal rigidity is achieved you can push the silver thumb tab on the valve system to lock the BeanBag into that position. Once that tab is depressed the white valve will “click” locked and no air will flow in or out of the BeanBag. 
At this point you can either leave suction applied as a secondary precaution or remove suction from the bladder as it is now “locked” close and will encounter air loss.

6. To reposition patient simply push the end of the white valve in and air will begin to flow back into the BeanBag.
*If patient is to be cradled by the BeanBag again you MUST apply suction so the BeanBag becomes rigid once again. 

!!Please make sure patient is supported by staff as the Beanbag will no longer support the patient!!
· Once all air is back into the relaxed Beanbag it can be cleaned, following current surface cleaning protocol established by your facility, and stored for the next procedure. This is a reusable product.

· Please store in a relaxed state on a flat surface where the BeanBag is not at risk of being punctured, dragged or torn.

· Please inspect and assess the functionality of this and any surgical device at regularly scheduled maintenance intervals as well as before each usage.

· Please familiarize yourself and staff with the functionality and usage of this and any surgical device before it is to be used on a patient.

*If product becomes compromised discontinue usage immediately.*
David Scott Company  59 Fountain St. Framingham, MA 01702  www.bdgel.com
508-875-3333  (fax) 508-875-3375 

- 1 –

Positioning Made Simple.doc

