Viamed oxygen sensor tests

Method

- Remove sensor from packaging, leave for 15 minutes to acclimatise before first use.

- Fit cell to vent, carry out normal set-up procedure. Repeat if any part fails.

- Run vent with a paed circuit and adult test lung on normal paed settings* for 10 to 15 mins at each of these O2 levels: 21%, 40%, 60% and 90%.

- Repeat that with a baby test lung on baby settings#.

- Soak test: leave running at 40% O2 for 2 hours or longer if time available.

- Fill in the form, noting anything unusual during the process.

- Replace the original O2 sensor and check the vent.

*Paed settings: PRVC, TV 500ml, resp rate 15, PEEP 8cm H2O, Ti 0.9s Trise 0.15s

#Baby settings: Pressure control, resp rate 30, pressures 25/5 cm H2O, Ti 0.5s, Trise 0.15s

Test 1 14/8/23

Cell s/n AA600018 opened 14:55 Test run began 15:45

Servo-i (universal)  PICU 12 s/n 115362

Test procedure: passed first time

Paed settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 41 | 61 | 91 |

Neo settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 22 | 40 | 61 | 91 |

Soak test: 1 hr. reading stable at 40%

At each setting, vent reached set reading in 15 to 20 seconds.

Test 2 15/8/23

Cell s/n AA600018 opened 14/8/23 Test run began 14:30

Servo-i (universal)  PICU 7 s/n 71613

Test procedure: passed first time.

Paed settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 40 | 60 | 90 |

Neo settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 40 | 58 | 91 |

Soak test: 15 mins at 40%, ran out of time. No problems.

At each setting, vent reached set reading in 15 to 20 seconds

Test 3  25/9/23

Cell s/n AA600017 opened 25/9/23 Test run began 14:30

Servo-i (universal)  PICU 10 s/n 13031

Test procedure: passed first time.

Paed settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 40 | 61 | 92 |

Neo settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 40 | 60 | 91 |

Soak test: 45 minutes at 40%, no problems.

At each setting, vent reached set reading in 20 seconds

Test 4 26/9/23

Cell s/n AA600018 opened 25/9/23 Test run began 14:30

Servo-i (adult)  PICU 14  s/n 165116

Settings: as above but Ti = 0.92, Trise =1

Test procedure: passed first time

Paed settings:

| O2 set: % | 21 | 40 | 60 | 90 |
| --- | --- | --- | --- | --- |
| O2 reads: % | 21 | 42 | 61 | 90 |

Soak test: 30 minutes at 40%, stable, no problems.

At each setting, vent reached set reading in 15 to 20 seconds