VM3COP40.13 R-33DE Oxygen Sensor Production Procedure

| Parts list |  |  |
| --- | --- | --- |
| Qty | Description | Part No. |
| 1 | Oxygen sensor – R-33S1 | 0110132 |
| 1 | 2 pin JST connector | 9071014 |
| 1 | Anti-static gas barrier bag | 0150000 |

Tool list

Soldering iron

Wire cutters

Stanley knife

Super glue

Small pliers

| Production: |  |
| --- | --- |
| - Remove the R-33S1 Oxygen sensor from the gas barrier bag. |  |
| - Using a soldering iron, remove the black and red wires. |  |
| - Ensure that the PCB is free from any excess solder or residue. |  |
| - Using the flat edge of a Stanley knife, remove a small area of surface lacquer from in front of the solder contacts. |  |
| - Using small pliers, remove the marked contacts. |  |
| - Splay the solder contacts to around a 30° angle. |  |
| - Place a drop of super glue on to the de-lacquered area of the PCB. |  |
| - Place the JST connector on to the super glue, ensuring that the connector contacts align with the contacts of the PCB. |  |
| - Solder the connector into place. |  |
| - Using a Stanley knife and steel rule, remove the excess from the bottom on the serial number label. |  |
| - Remove the Teledyne label (R-33S1) and replace it with the Vandagraph (R-33DE) label with the same serial number. NB. To ensure retention of the original serial number, each sensor should be done individually. |  |
| QA: |  |
| - Using the appropriate test lead, connect the sensor to a digital volt meter. |  |
| - Set the digital volt meter to the millivolt scale. - Observe the output of the sensor. The output must be between 23mV and 27mV. - Using the Intrastats system, mark the sensor as having passed or failed QA as applicable. |  |
| Packing: |  |
| - Take a sensor that is to be packed, note its serial number, and affix that corresponding bag serial number label to the bag. |  |
| - Place the sensor in the bag. - Using a strip heat sealer, seal the sensor into the gas barrier bag. - Once passed, place the finished product, along with the production sheet, on to the production shelf to await stock entry. |  |