Observations on Maxtec training presentation for MaxBlend 2
Steve Hardaker - 17/6/24

Observations regarding Maxtec’s first draft training presentation for MaxBlend 2

P2 – Viamed MaxBlend2 Training – should this be Maxtec MaxBlend2 Training so it appears authorised by Maxtec instead of Viamed?

P3 – 10 seem to be corporate promotion, might be best left out of this.

P11 – Picture of an incubator, would be better with a picture of the MaxBlend2

P18 - 4. Calibrate the sensor prior to use according to the calibration procedures in section 2.8 [of the Instructions For Use] and as detailed further into the training presentation (P24).

P20 – Needs a step 3 to say how to set the alarm limit, suggest copying this from the manual:

NOTE: The Low Alarm digits begin to flash indicating the Low Alarm manual setting.
Use the UP and DOWN keys to set the low alarm to the desired value.
Pressing the arrow keys changes the value in 1% increments. If the keys are held down
for more than 1 second, the display will scroll at a rate of 1% per second.

P21 - Same for High alarm.

P26-27 – References to 50PSI. I think this section is irrelevant to end user training as it is for engineers only.

I don’t see any harm leaving P31 – 34 Service & Maintenance in there, the end users will likely not need this but it gives them some basic troubleshooting and testing guidance.

P37 – Viamed’s contact details and lose the stormtrooper! :-)

Additions:

It needs an explanation of the toggle and when to use it.

A statement regarding if the performance of the device is in question, to contact the hospital’s Medical Engineering department or persons responsible for equipment maintenance.

A brief summary of the gas alarms, to advise that the alarm will sound if the gas supply pressures are out of balance, for example if being operated from a cylinder that is becoming depleted.

It doesn’t actually state anywhere that the oxygen concentration can be adjusted from 21% to 100% using the knob on the front. - From the manual: The oxygen concentration knob allows the clinician to select a desired oxygen concentration from 21% to 100% O2.