The VN202 mkII has a wide calibration span and therefore can be used as a percentage oxygen analyser over a range of atmospheric pressures. As a result you can ‘normalize’ the air calibration to 20.9% at altitude. If the VN202 mkII is used in the advised mode as a percentage oxygen analyser - ensure that you calibrate every time the analyser is used; and if variance occurs in local atmospheric pressures (due to change in altitude or weather conditions).

However, there may be occasions when the VN202 cannot adjusted to an air calibration value of 20.9%, and therefore cannot be used as a percentage oxygen analyser:

a) At extreme altitudes.

      and/or

b) Because of output limitations of the sensor due to: age of sensor, ambient temperature…

If this occurs you can use the analyser as a partial pressure analyser. To do this you will need to make some calculations - using known barometric pressure, or approximations using a compensation table.
