	   [image: image3.jpg]



	VIAMED[image: image1.png]





Post Market Surveillance Review to 21st  March 2024
4310002, 4310009 Flowsensor SpiroTrue A (PC). Box of 5 
Bluepoint Medical references: 7030131001, 3030131003 
First sold: 24th May 2013 
Total sales to date: 13,435 boxes, 
                                91 different customers across 30 countries 
Total failures: 10 boxes 
Total returns: 10 boxes 
Total customer failures: 0.074% 
Royal Brompton Hospital SRS64245, received 1 box 19th March 2014 
Inaccurate tidal volume readings. Sometimes on 500ml inspired volume, the A(PC) will read up to 590ml, other times just a little out, e.g. +10ml. It is not an issue with the ventilators, as the problems do not exist when they revert back to original Draeger flowsensors. Customer originally purchased Flowsensor SpiroTrue A, but switched to SpiroTrue A (PC) due to MHRA stance on single use designation. 
Result: Customer stopped purchasing, their requirements is about 60 boxes per year. Will try and regain the sales, with either Flowsensor SpiroTrue A or Flowsensor SpiroTrue A (PC). 
James Paget Hospital SRS64307, received 5 boxes 16th May 2014 
Calibration issues when used with Evita 4. Customer originally purchase Flowsensor SpiroTrue A, but switched to SpiroTrue A (PC) due to MHRA stance on single use designation. 
Result: Customer retained.
Royal Bournemouth SRS64386, received 4 boxes 15th July 2014 
Calibration issues on Draeger Evita XL, this happened on six different ventilators. Some failed initial calibration, other failed shortly afterwards. Customer originally purchase Flowsensor SpiroTrue A, but switched to SpiroTrue A (PC) due to MHRA stance on single use designation. 
Result: Customer stopped purchasing, their requirements is about 15 boxes per year. Will try and regain the sales, with either Flowsensor SpiroTrue A or Flowsensor SpiroTrue A (PC). 
For the period from 16th July 2014 to 21st March 2024, there has not been any new reported failures or failures that refer to the claims listed above. 
The design, functionality and reliability of the sensor is very good. We do include feedback forms within each box of sensors for end users to be able to easily report possible issues and/or for confirming compatibility. Unfortunately, we get very few feedback forms returned as the customers are happy with the sensors and have nothing to report.
Applications: 
	OEM
	OEM Part Number
	Equipment Type

	Acutronic
	Evita range
	Ventilator

	Acutronic
	Savina range
	Ventilator

	Asmuth GmbH Medinzintechnik
	2 42516796 0167 2
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	6030131004
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	AS432+-AUTO
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	AS434
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	SpiroTrue APC
	Flow Sensor

	Bluepoint Medical
	2 42516796 0157 7
	Flow Sensor

	Bluepoint Medical
	3030131003 
(sold 1 piece per box)
	Flow Sensor

	Bluepoint Medical
	SpiroTrue APC
	Flow Sensor

	Draeger
	Apollo
	Anaesthesia System

	Draeger
	AV1
	Anaesthesia System

	Draeger
	Cato
	Anaesthesia System

	Draeger
	Cicero
	Anaesthesia System

	Draeger
	D-9119-2009
	Flow Sensor

	Draeger
	Dura
	Ventilator

	Draeger
	Eva
	Ventilator

	Draeger
	Evita
	Ventilator

	Draeger
	Evita 4
	Ventilator

	Draeger
	Evita XL
	Ventilator

	Draeger
	Fabius GS
	Anaesthesia System

	Draeger
	Fabius range
	Anaesthesia System

	Draeger
	Julian
	Anaesthesia System

	Draeger
	MK01900 
(autoclavable version of Spirolog)
	Flow Sensor

	Draeger
	Perseus A500
	Anaesthesia System

	Draeger
	PM 8030
	Monitor

	Draeger
	PM 8035
	Monitor

	Draeger
	PM 8050
	Monitor

	Draeger
	PM 8060
	Monitor

	Draeger
	Primus
	Anaesthesia System

	Draeger
	Primus / Apollo
	Anaesthesia System

	Draeger
	Savina
	Ventilator

	Draeger
	SpiroLife 
(autoclavable version of Spirolog)
	Flow Sensor

	Draeger
	SpiroLife MK 01900
	Flow Sensor

	Draeger
	SpiroLife MK01900
	Flow Sensor

	Draeger
	Sulla
	Anaesthesia System

	Draeger
	Zeus
	Anaesthesia System

	Viamed
	2 42516796 0168 9
	Flow Sensor

	Viamed
	7030131001
	Flow Sensor

	Viamed
	Flowsensor A (PC)
	Flow Sensor

	Viamed
	SpiroTrue A (PC)
	Flow Sensor


Reviewed by: Steve Nixon
D[image: image2.png]Autoclavable



irector – Viamed Ltd.
Date:  21st  March 2024
