	   [image: image3.jpg]



	[image: image1.png]



VIAMED


Post Market Surveillance Review to 21st March 2024
4310001, 4310007 Flowsensor SpiroTrue A. Box of 5 
Bluepoint Medical references.: 1030132006, 3030131004
First sold: 12th January 2009 
General sales suspended 18th December 2013 due to MHRA stance on single use designation. Reinstated when covered by Bluepoint Medical CE certification. 
Total sales to date: 5,560 boxes, 
                              120 different customers across 34 countries 
Total failures: 6 boxes
Total returns: 6 boxes
Total customer failures: 0.107%
Liverpool Heart & Chest Hospital SRS63066, received 1 box 10th October 2011 
Failed calibration. Also reported sensors generally have shorter life than Draeger OEM sensors. 
Result: Customer lost, approximately 100 boxes per year. Resumed using Draeger OEM sensors, will try and regain the sales with Flowsensor SpiroTrue A.
Dorset County Hospital SRS63869, received 1 box 3rd June 2013 
Failed calibration. 
Result: Customer lost, approximately 12 boxes per year. Will try and regain the sales with Flowsensor SpiroTrue A.
Papworth Hospital SRS64002, received 4 boxes 11th September 2013 
Failed calibration when used on Evita 4s. 
Result: Customer lost, approximately 120 boxes per year. They have tried both SpiroTrue A and SpiroTrue A (PC). Will try and regain the sales with Flowsensor SpiroTrue A.
For the period from 12th September 2013 to 31st  July 2022, there has not been any new reported failures or failures that refer to the claims listed above. 
The design, functionality and reliability of the sensor is very good. We do include feedback forms within each box of sensors for end users to be able to easily report possible issues and/or for confirming compatibility. Unfortunately, we get very few feedback forms returned as the customers are happy with the sensors and have nothing to report.
Applications: 
	OEM
	OEM Part Number
	Equipment Type

	Acutronic
	1601
	Flow Sensor

	Acutronic
	Evita range
	Anaesthesia System

	Acutronic
	Savina range
	Ventilator

	Acutronic
	SpiroQuant A
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	2 42516796 0158 0
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	2030131000
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	AS432+
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	AS8403735
	Flow Sensor

	Asmuth GmbH Medinzintechnik
	SpiroTrue A
	Flow Sensor

	Biometric Cables
	910246001
	Flow Sensor

	Biometric Cables
	FS1
	Flow Sensor

	Bluepoint Medical
	1030132006
	Flow Sensor

	Bluepoint Medical
	2 42516796 0157 3
	Flow Sensor

	Bluepoint Medical
	6871980 
	Flow Sensor

	Bluepoint Medical
	PM 8030
	Monitor

	Bluepoint Medical
	SpiroTrue A
	Flow Sensor

	Draeger
	Spirolog 84 03 735
	Flow Sensor

	Draeger
	84 03 735
	Flow Sensor

	Draeger
	8403735
	Flow Sensor

	Draeger
	Anemone
	Monitor

	Draeger
	Apollo
	Anaesthesia System

	Draeger
	AV1
	Anaesthesia System

	Draeger
	Cato
	Anaesthesia System

	Draeger
	Cicero
	Anaesthesia System

	Draeger
	D-23802-2015
	Flow Sensor

	Draeger
	Dura
	Ventilator

	Draeger
	Eva
	Ventilator

	Draeger
	Evita
	Ventilator

	Draeger
	Evita 4
	Ventilator

	Draeger
	Evita XL
	Ventilator

	Draeger
	Fabius GS
	Anaesthesia System

	Draeger
	Fabius range
	Anaesthesia System

	Draeger
	Infinity ID
	Flow Sensor

	Draeger
	Infinity V500 
	Anaesthesia System

	Draeger
	Julian
	Anaesthesia System

	Draeger
	MT-2086-2007 
	Flow Sensor

	Draeger
	PM 8030
	Monitor

	Draeger
	PM 8035
	Monitor

	Draeger
	PM 8050
	Monitor

	Draeger
	PM 8060
	Monitor

	Draeger
	Primus
	Anaesthesia System

	Draeger
	Savina
	Ventilator

	Draeger
	Spirolog
	Flow Sensor

	Draeger
	Spirolog 8403735
	Flow Sensor

	Draeger
	Sulla
	Anaesthesia System

	Envitec
	1000470
	Flow Sensor

	Envitec
	SpiroQuant A+
	Flow Sensor

	Integral Process
	0032 00 02
	Flow Sensor

	Integral Process
	65025
	Flow Sensor

	Integral Process
	8403735
	Flow Sensor

	Intersurgical
	2777000
	Flow Sensor

	Intersurgical
	e-FlowSensor
	Flow Sensor

	Lub Dub Medical
	OX-2529
	Flow Sensor

	Medacx
	M-45-00-0005
	Flow Sensor

	Medacx
	SpiroQuant A
	Flow Sensor

	Proact Medical
	1000470
	Flow Sensor

	Proact Medical
	SpiroQuant A+
	Flow Sensor

	Viamed
	2 42516796 0160 3
	Flow Sensor

	Viamed
	3030131004
	Flow Sensor

	Viamed
	Flowsensor A
	Flow Sensor

	Viamed
	SpiroTrue A
	Flow Sensor


Reviewed by: Steve Nixon
Director – Viamed Ltd.
[image: image2.png]



Date: 21st  March 2024
