[image: image1.png]N

[ E )= &)

Sleeve lock

Securing segment
End device

Cover cap

Internal electrical line
Slot

Sleeve



[image: image2.png]‘11,5 mm \
Y

min. 1,5 mm
Dicke
Thickness

Securing segment





Nufer Medical AG

Telefon
0041 (0)31 958 66 66       
Schürmattstrasse 6

Fax
0041 (0)31 951 46 73
Postfach 125

Email
info@nufer-medical.ch

CH-3073 Gümligen
Internet
www.nufer-medical.ch
Schweiz
MWST-Nr. 
121 519

Nufer Medical  -  let’s find the best solutions


INSPECTIONS AND MAINTENANCE ADVICE
(Supplement to the operating instructions)

CERATHERM 600-2
WALL, CEILING AND MOBILE MOUNTING UNITS
· Yearly inspection: The wall, ceiling and mobile mounting device for the Ceratherm 600-2 must be inspected once a year by the operator with regard to:

1.
 Damage to paintwork
2.
 Cracks on plastic parts

3.
 Deformation of the pendant system

4.
 Loose parts

In case of any faults or damage occurring, please contact Nufer Medical AG.
·  Check of securing segment every two years: A qualified technician must check and grease the securing segment at least every two years as follows:

1.   Remove end device (C).
2. Check securing segment (B) for a minimum thickness of 1.5 mm and minimum length of 11.5 mm and replace as required (Spare part number 521-37247).
3. Grease securing segment (B) and pin of end device (C) with Dow Corning Molykote PG54 Silicone Grease.
4. Install end device (C) and make a functional test.
[image: image3.png]



[image: image4.jpg]r

NUFER-MEDICAL





For any further assistance or spare parts, do not hesitate to contact Nufer Medical AG.





















18.002.B_INSPECTIONS_MOUNTING_UNITS_EN.doc                       Nufer Medical AG / 09.12.2008 / DU


Seite 1/1

