[NHS Trust logo removed]

Training Delivery Plan - MAX O2+AE

| Section | Information |
| --- | --- |
| Introduction | The Max02+AE is a type of oxygen analyser used to spot check or measure the concentration of oxygen being delivered to patients ranging from newborns to adults. It is important the device is used as intended and appropriately to prevent the cause of inaccurate oxygen readings which can result in improper treatment, hypoxia and hyperoxia. |
| Considerations & Contraindications | General EME asset label / service due date Inspect the device for any visible damage, such as cracked screen or visible hairline cracks. If any damage is found or device is not functioning, then contact EME. Contraindications Do not use for continuous monitoring of oxygen delivery to a patient. The device does not have alarms so is only effective when observed by the user. Precautions: Patient Do not install the tee adaptor in a location that will expose the tee adaptor to patient’s exhaled breath or secretions. Never allow the sensor cable near the patient’s head or neck as it may result in strangulations. Precautions: Max02+AE Calibrate the Max02+AE weekly when in operation, or if environmental conditions change significantly. Do not use in an MRI environment. If the device is exposed to liquids (spills or immersion), turn OFF and then ON. This will allow the unit to self-test to ensure the device is operating correctly. Store the MaxO2+ analyser in a temperature similar to its ambient environment of daily use. |
| Description of device appearance | The Max02+AE consist of a base that has the following features (show and explain): - 3-Digit LCD display: displays that the percentage (%) of the oxygen concentration. - Low battery indicator: the low battery indicator will only be displayed when battery level is below the normal operating level. Contact EME to replace the batteries. - Calibration Key: used to calibrate the device, holding the key for more than three seconds will allow the device to enter into calibration mode. - On/Off Key: to switch on/off the device Also, the device consists of a sensor cable and sensor attached to the cable (show). |
| Replacement Parts and Accessories | Max-250E Oxygen Sensor attached to the sensor cable would be replaced yearly by the EME team as part of yearly maintenance. Blue tee Adapter (show) used to connect the Max02+AE to the machine via the patient circuit. Precaution: The Blue Tee Adapter must be replaced in event of contamination i.e. exposed to patient’s exhaled breath or secretions. The replacement can be obtained from the Resus cupboard or from Medical Equipment Library. |
| Demonstrate how to calibrate of the device | The device automatically calibrates to the room air after switching on the device each time. The Max02+AE needs to be calibrated weekly, as a reminder, one-week timer is started with each new calibration where its ‘CAL’ will be displayed. - Press the calibration key for more than 3 seconds until ‘CAL’ displays on the screen - Hold the sensor in the air until the calibration gas of 20.9% (room air) is displayed on the screen. Record the date of calibration of the device on the NIV checklist (show). |
| Demonstrate how the device is connected to the Trilogy Evo: | Attach one side of the Blue Tee Adaptor to the outlet of the Trilogy Evo device and the other side to the filter of the patient circuit (show). (Ensure the tee adaptor is attached before the filter to prevent contamination) Insert the Max02+AE in the centre position of the tee adaptor (ensure the Max02+AE is secured in the tee adaptor). |
| Monitoring of the Fi02 level (%). | Hourly spot check should be conducted, where the device can be left switched on continuously. Monitor the Max02+AE device and verify the Fi02 level (%) set on the Trilogy device matches or is within the tolerance of 5% displayed on Max02+AE. |
| Cleaning the equipment | Base unit and the Tee adaptor Use a mild detergent and a moist cloth (Clinell wipes -as per Trust guidance). Take appropriate care to prevent any solution from entering the instrument. DO not immerse unit in fluids. Tee adaptors must be thrown away in event of contamination and replace it with new Tee adaptors. Oxygen sensor Use a cloth moistened with isopropyl alcohol (sani- cloth 65% alcohol/water solution). Do not throw the sensor away. |
| Storage of device | One device is permanently situated in the Resuscitation cupboard, and if more required, please contact the Medical Equipment Library (MEL): QEQM 725-3892 |
| Issue with or failure of device | - “Cal Err St” may occur when the sensor has not reached thermal equilibrium (normal operating temperature) - Allow adequate time for the sensor to equilibrate to a new ambient temperature - “Cal Err lo” may occur when the oxygen sensor requires changing - Contact and return the device to EME for the oxygen sensor to replaced. - “EO4” occur when the battery is below minimum operating voltage - Contact and return the device to EME for batteries to replaced. |
| User manual location | Contact your ward superuser to get access to the user manual. |