Viamed Management Review
2024

Written by:

Derek Lamb
Managing Director

Minutes by: Helen Lamb 4th November 2024

### Period September 2023 – October 2024

# (0.0) VIAMED Board ISO Review Directors Meeting

Objective: Should be held at least yearly or as required
### Report References
Ref #343435.14663 Current Objectives Pre Meeting
Ref #336857.14666 ISO 13485 Scope for review
Ref #336857.14667 Exclusions To 13485 Review
### Notes
Must review and sign off the Scope and Exclusions
Last meeting held 18th October 2023

Jean Lamb, Steve Nixon Helen Lamb and Derek Lamb all present. Minutes taken by Helen lamb. Meeting started at 10.15am
Agenda and objectives was reviewed by the board before the start of the meeting and agreed upon.

Reviewed Viamed Scope does not say MDD. Agreed to drop first sentence of scope service of Supra-Maximal nerve simulator, simulation equipment etc.

Just leave Distribution and Service of Medical Devices and Products. Derek will discuss with BSI scheme manager.
Reviewed the current exclusions on the 13485 standard. All happy with the current exclusions.
Reviewed the current Viamed Quality Policy all agreed happy with it.
ISO has been in and we have one non conformance, we were missing a part of our procedure on calibration. Need a date label on each unit. So we need to update the VOP to include this and add expire date to each piece of equipment. We will complete the non conformance and return to BSI this week.

### Follow on Issues

Management review Board Meeting will be held in 12 to 18 months from now.
345784 issue sent to go to scheme manager to update the scope of our 13485 certificate.
### Conclusion

 Board meeting held today, all present.

# (1.0) Minutes of Last Meeting
Objective: Produce minutes of previous meeting, is everyone happy with them.
### Report References
Ref #336857.14695 Minutes of Last Meeting
### Notes
  Meeting Date: 18 October 2023
- Attendees:
  - Jean Lamb
  - Steve Nixon
  - Derek Lamb
  - Helen Lamb (Minutes taker)
- Meeting Start Time: 10:30 AM
- Meeting End Time: 5:00 PM
- Agenda: Reviewed and agreed upon by the Board before the meeting started.

Summary:
The 2023 Management Review Board meeting was conducted on 18 October 2023. Attendees included Jean Lamb, Steve Nixon, Derek Lamb, and Helen Lamb, with the latter responsible for taking the minutes. The meeting lasted from 10:30 AM to 5:00 PM. The agenda was reviewed and approved prior to the meeting.

All reviewed the minutes of the last Management Review Board Meeting and agreed all happy.

### Follow on Issues
###
### Conclusion
No further action required.

# (2.0) Matters Arising
Objective: To ensure actions of previous meeting have been actioned.
### Report References
 302478.13321 Viamed 2023 Management Review Minutes included Oct 23 - Helen Lamb Minutes_10_10_24.pdf
Issue 277097 still ongoing and being worked on. - Still ongoing ok to be left.
Issue 277266 still ongoing, Distributor agreement is being rewritten to utilise the new system. - Issue closed but still working on
Issue 309724 was sent as there is a system bug. - Done Issue closed.
Issue 309727 sent to DL to set the objective for 2024. - Done Issue closed.
Issue 309728 to DL regarding making a report that shows accounts on stop and the total value. - Done Issue closed.
Issue 309736 sent to DL regarding reviewing documents per supplier, to be discussed in the Wednesday meeting. - Done Issue closed.
Issue 309737 sent to DL to monitor/review out-of-target stock. - Done Issue closed.
Issue 309746 sent to DL to test that the new issue, Adverse incidents, shows on the review reports. - Done Issue closed.
Issue 309748 sent to set up a system to get feedback sent to suppliers. - Issue still ongoing more needs to be done for this.
Issue 309770 sent to DL for moving the Process Update Effectiveness Review from this section to section 13.3 Process Performance. - Done Issue closed.
Issue 309983 sent to HL for follow-up on building fabric matters. - Done Issue closed.
Issue 309984 sent to all staff regarding Coronation bank holiday addition to the 2024 holiday allocation. - Done Issue closed.
Issue 309857 sent to SN to explore UKCA and UKRP training for SH. - Still ongoing for further discussion. Maybe someone else does this. Issue updated to reflect this.

### Notes
Note above on bold to reflex status of the issues.
### Follow on Issues

### Conclusion
No further action required.
# (2.1) Follow-up actions from previous reviews

Objective: To ensure previous reviews are addressed and resolved
### Report References

### Notes
SEE (2.0)	Matters Arising Done together.
Header: (2.0) Matters Arising
Follow-on Issue:
- Issue 277097 still ongoing and being worked on.
  Notes : (Export sales were dropping, - it recovered a bit and is now between the 2 years)
- Issue 277266 still ongoing, Distributor agreement is being rewritten to utilize the new system.
  Notes : now Closed

Header: (3.0) Turnover and Predicted for Year
Follow-on Issue:
Issue 309724 was sent as there is a system bug.
Notes : 07 Dec 2023 Derek Lamb updated the turnover pages, stock sales pages, and country sales so far,couple more have been done in the stock groups pages,rest will do as and when they sow themselves,not all existing pages were affected

Header: (6.0) Target for Year
Follow-on Issue:
Issue 309727 sent to DL to set the objective for 2024.
21 Oct 2023 Derek Lamb done

Header: (10.0) Debtors
Follow-on Issue:
Issue 309728 to DL regarding making a report that shows accounts on stop and the total value.
17 May 2024 Derek Lamb accounts on stop shows on the front crm page,unable to show value as that's in the accounts xero packages

Header: (11.1) Purchasing controls
Follow-on Issue:
Issue 309736 sent to DL regarding reviewing documents per supplier, to be discussed in the Wednesday meeting.

Header: (13.0) Stock Levels and product performance
Follow-on Issue:
Issue 309737 sent to DL to monitor/review out-of-target stock.
26 Jan 2024 Helen Lamb this has now come down

Header: (15.0) Customer Complaints
Follow-on Issue:
Issue 309746 sent to DL to test that the new issue, Adverse incidents, shows on the review reports.
21 Oct 2023 Derek Lamb 31015721 Oct 2023 test issue works as shown

Header: (15.1) Customer and product feedback
Follow-on Issue:
Issue 309748 sent to set up a system to get feedback sent to suppliers.
21 Oct 2023 Still a work in progress, survey monkey customer feedback working, data going into intrastats, lots currently changing with regardd suppleirs, data reviews and climate Questions

Header: (15.4) Preventive and corrective actions
Follow-on Issue:
Issue 309770 sent to DL for moving the Process Update Effectiveness Review from this section to section 13.3 Process Performance.
17 May 2024 Derek Lamb
done

Header: (18.0) Building fabric Issues
Follow-on Issue:
Issue 309983 sent to HL for follow-up on building fabric matters.
13 Jun 2024 Helen Lamb
this has all been done. We have also now got a gardener and had other work completed

Header: (19.0) Responsibilities, Staff, Training Issues
Follow-on Issue:
Issue 309984 sent to all staff regarding Coronation bank holiday addition to the 2024 holiday allocation.

Header: (19.1) Staff Appraisals
Follow-on Issue:
Issue 309857 sent to SN to explore UKCA and UKRP training for SH.

### Follow on Issues

### Conclusion

No further action required.
# (3.0) Turnover and Predicted for Year
Objective: 3.0 Million
### Report References

Ref #336857.14643 Turnover Report

### Notes
Summary
The turnover data from 2022 through 2024 shows fluctuating trends.
In 2023, the total turnover was £2.669 million, and in 2022, it was slightly lower at £2.547 million.
For 2024, the projected turnover is £2.63 million, based on the performance from January through October (£2.037 million) and estimated values for the remaining months global.
However, the current figures for 2024 fall short of the £3.0 million objective, with significant variances in the second half of the year, particularly in October. Additional effort may be required to meet the goal.

UK is about the same as last year. Export no Eu and EU all very similar to last year. All a little down.
Automotive is slightly up.
### Follow on Issues

### Conclusion

 On target for next year objective remaining same.

# (4.0) Profitability

Objective: 5 Percent Net of Turnover
### Report References

### Notes
Heading: Profitability
Objective: 5 Percent Net of Turnover
We need to be hitting 120,000 to 130,000 for 5% of gross profit.
Currently we been running at 1% at best for 2023.
Looks like we are running at 1%. Prices to be reviewed but no overall increase this year.
Getting a live profit and lose with help from accountant. So we hope to be getting a more accurate state of current.
Every month the system is counting barcode so we have more accurate and live stock figure. 2025 we should be able to have a live view of where we are at any one time.
### Follow on Issues

### Conclusion
We are not on target but we are working towards an accurate figure showing here we are at any time.

# (6.0) Target for Year
Objective: Really need to be Aiming for 3 Million.	To be sure to be profitable.
### Report References

### Notes

Still needing to be aiming for 3 million currently at 2.65 million.

### Follow on Issues

### Conclusion
Off target.

# (9.0) Overdraft
Objective: Target staying out of using the overdraft

### Report References

### Notes
 We currently have no overdraft.

### Follow on Issues

### Conclusion
No overdraft and no expectation at this time of needing one.
On target.

# (10.0) Debtors
Objective: To be below 15 Percent of turnover

### Report References

Notes
Raw Data:
- Invoices Owed to You:
  - 592 Awaiting payment (including overdue): £260,657.59
  - 229 Overdue invoices (subset of above): £131,726.92
- Invoices by Timeframe:
  - Older: Significant volume.
  - 29 Sep–5 Oct: Low volume.
  - This week: Average volume.
  - 13–19 Oct: Increasing volume.
  - 20–26 Oct: Moderate volume.
  - Future: High volume.
- Turnover (Yearly):
  - 2024 Estimated Turnover: £2,037,263 (current) | £2,630,419 (estimated for the full year)
### Concise Summary:
- Total invoices awaiting payment: £260,657.59, of which £131,726.92 (50.5%) are overdue.
- The company aims to keep total debtors below 15% of turnover. Based on the current turnover for 2024 of £2,037,263, 15% equals approximately £305,589.
- Current total debt of £260,657.59 is within the 15% threshold, representing 12.8% of the current turnover.
- However, 50.5% of the total outstanding amount (£131,726.92) is overdue, which is a significant proportion.
### Action Points:
1. Prioritize the collection of overdue invoices to reduce the risk of further overdue debt, especially the £131,726.92.
2. Monitor the total debtor balance as turnover progresses to ensure the company remains under the 15% threshold for the full year.
Receivables breakdown by notable entries:
- Barts Health NHS Trust: £7,745.04 due in <1 month.
- D. Cortessis & Co L.P.: Total of £15,179.35, of which £6,391.64 and £8,787.71 in 3 months (FX).
- JJS Manufacturing Ltd: £19,137.60 total, with a significant portion under current, <1 month, and 1 month.
- Premier Diagnostics Ltd: £7,582.20 total, primarily older.
- Oxylink Piotr Podkowa (Euro): £15,904.75 total, primarily 1 month and older.
### Overall Total and Summary
The total outstanding receivables across all periods for Viamed Limited as of 31 October 2024 is £321,379.99.
### Breakdown by Ageing Categories:
- Current: £39,679.17
- <1 Month: £112,763.18
- 1 Month: £38,026.73
- 2 Months: £12,002.98
- 3 Months: £16,826.60
- Older: £102,081.33

### Follow on Issues

### Conclusion
All happy, no change to objective, happy with current status of debtors.

# (10.1) Contract review , Picking, Packing and Dispatch
Objective: Review Contract review, Picking, Packing and Dispatch

### Report References
Ref #336857.14644 Reasons for Editing Customer Orders
Ref #336857.14645 Shipping Stats
Ref #336857.14646 Shippers used
Notes
 Ref #336857.14644 Reasons for Editing Customer Orders
Key Takeaways:
- Customer Order Changes and Customer Updates are the most frequent reasons for order edits in the "Customer Update" category.
- Input Errors, such as Incorrect Information and Address Amendments, peak in the early months of 2024.
- Layout Errors remain infrequent but occur sporadically, with no major trends.
Regular business updates, like Priority and Price Changes, happen mostly during early 2024, with some consistency in carriage and shipment-related changes.
Ref #336857.14645 Shipping Stats
### Summary:
- Total Orders: 5,272 orders recorded from July 2023 to September 2024.
- Breakdown by Shipping Time:
  - Same Day: 3,260 orders (61.83%)
  - Next Day: 914 orders (17.33%)
  - Combined Same or Next Day: 4,174 orders (79.16%)
  - 3 Days: 153 orders (2.9%)
  - 4 Days: 193 orders (3.66%)
  - 5 Days: 90 orders (1.7%)
  - 6 Days: 64 orders (1.21%)
  - 7 Days: 52 orders (0.98%)
  - More than 7 Days: 544 orders (10.31%)
- Key Observations:
  - 79.16% of orders are shipped within the same or next day.
  - A significant portion of delays occurs in orders taking more than seven days to ship (10.31%).
  - There is a noticeable spike in orders shipped on the fourth day (likely influenced by weekend delivery gaps).
- Notes:
  - Data is based on approximations and considers the time between order confirmation and shipment notifications.
  - Orders processed over the weekend may reflect in 4-day shipments, causing minor spikes.
This summary highlights that nearly 80% of orders are shipped within one day, indicating strong performance in quick deliveries.

Ref #336857.14646 Shippers used
Summary:
1. Overall Volume: The total parcel volume ranged between 171 and 444 monthly deliveries from May 2023 to October 2024. The highest volume was in April 2024 with 444 parcels, while the lowest was in October 2024 with 171 parcels.
2. Key Courier Services:
  1. UPS: Saw consistently high usage, peaking in September 2024 with 405 parcels.
  2. DHL: Frequently used, ranging between 1 and 7 parcels monthly.
  3. FedEx: Generally managed between 1 and 6 parcels, with an increased usage in June 2024 (5 parcels).
3. Variations: Usage of less common couriers like EMS, Parcel Force, and Kuehne Nagel fluctuated with little consistency.
4. Trend: There was a noticeable increase in parcel volume in April 2024, followed by a gradual decline in subsequent months.
Goods out we now have Emma doing double hours so we are now more effective goods out.
### Follow on Issues

### Conclusion
We have reviewed Contract review, Picking, Packing and Dispatch and we are all happy with this.
All happy with the competency of office and warehouse staff, and we find them to be effective.
# (11.0) Creditors
Objective: To have no overdue Creditors
### Report References
 Ref #336857.14649 Creditors Report

### Notes

### Follow on Issues

### Conclusion
Overdue are council tax bills, Viamed Properties has two rental bills outstanding, but the funds are low so it has been decided to leave these until funds are more available. No other overdue creditors.
No issue with creditors.
All happy with creditors.

# (11.1) Purchasing controls
Objective: Review Purchasing controls, is everyone happy with this.

### Report References
 Ref #336857.14650 Supplier Review Index
Notes
   Summary:
- Total Suppliers Reviewed: 36
- Suppliers with Rating "A" (Critical, sole supplier): 3 (e.g., H-Squared UK Electric Ltd.)
- Suppliers with Rating "B" (Supplier performing well): 31
- Suppliers with Rating "C" (Supplier OK, alternatives worth considering): 2 (The West Group LTD, Unimed)
Key Notes:
- Majority of the suppliers are rated B, indicating satisfaction but no exclusivity.
- Several suppliers lack formal supply agreements, but have ISO certifications, ensuring standards compliance.
- Few suppliers are not actively used or have expired reviews, requiring updates (e.g., Unimed, Beijing Choice Electronic).
Orders to Suppliers now cannot proceed unless there is a current and up to date supplier review been signed off.

We need to expand supplier review, we need a lot more up the chain information and reports. Including human slavery, carbon reduction plans, RoHS and REACH declarations, conflict minerals etc. New system is required.
Derek to add a column on the supplier documentation to include confirm documents still valid. The new request for information can be initiated from here. Adding headings for each new document required.
### Follow on Issues
345816 issue sent to update document index to facilitate getting more supplier documents and to verify existing ones more effectively. With the aim to satisfy the NHS requirements regarding Scope 3 up stream verification of suppliers.

### Conclusion
All happy with purchasing controls. All happy with current purchase order processing.
Supplier reports were effective over the last 13 months, moving forwards we require more information and documentations.

# (12.0) Loans
Objective: Review any loans and check all agree happy with current
### Report References

### Notes
 We don’t have any loans currently

### Follow on Issues

### Conclusion
Reviewed, no loans and all agree happy with current.

# (13.0) Stock Levels and product performance
Objective: Should not Exceed 25 Percent of Turnover
### Report References
Ref #336857.14651 Stock Movements
Ref #336857.14652 Product Fail reviews
Ref #336857.14653 QA Fails Report
Notes
Below is a detailed list of issues requiring action, excluding those marked "No Work Required."
Issue ID	Date Created	Completed By	Date Completed	Repair Details	Days to Complete
342882	04 Oct 2024	Derek Lamb	10 Oct 2024	0110043: Requires Cleaning! - Manufacturing Defect; 0310265: Requires Cleaning! - Full Service, different model/part number, Waiting on response from Envitec.	7 days
340294	06 Sep 2024	Derek Lamb	13 Sep 2024	0110215: Requires Cleaning! - Zero offset out of tolerance, recalibration; 7910110: Requires Cleaning! - On/Off button fault, no cap required.	9 days
338382	16 Aug 2024	Derek Lamb	05 Sep 2024	8010004: Requires Cleaning! - Connector unstable, VST sensors, no cap required at this time.	22 days
335077	12 Jul 2024	Derek Lamb	18 Jul 2024	4410530: Requires Cleaning! - Calibration test, no issues found.	8 days
333683	28 Jun 2024	Derek Lamb	04 Jul 2024	0110043: Requires Cleaning! - Sensor mechanically broken; 7910110: Requires Cleaning! - No fault found, no cap required.	7 days
333078	21 Jun 2024	Derek Lamb	04 Jul 2024	7910201: Requires Cleaning! - Not switching on, no cap required, Vandagraph Tek Ox.	15 days
332460	14 Jun 2024	Derek Lamb	20 Jun 2024	1114007: Requires Cleaning! - Credited stock, return to shelf, no cap required.	8 days
331778	07 Jun 2024	Derek Lamb	13 Jun 2024	0110579: Requires Cleaning! - No fault found.	7 days
331099	31 May 2024	Derek Lamb	06 Jun 2024	7910201: Requires Cleaning! - Intermittent power, Tek Ox Vandagraph, no cap required.	7 days
329856	17 May 2024	Derek Lamb	21 May 2024	0110217: Requires Cleaning! - Service replacement; 0380030: Requires Cleaning! - Serviced/adjusted factory set valves; 2810004: Requires Cleaning! - Display fault; 7910110: Requires Cleaning! - Sensor lead failed.	5 days
327633	26 Apr 2024	Derek Lamb	03 May 2024	971023: Requires Cleaning! - Multiple faults, sensor replaced, case damaged, replaced.	9 days
325169	02 Apr 2024	Derek Lamb	05 Apr 2024	7910201: Requires Cleaning! - Erratic reading; 9710203: Requires Cleaning! - Case fault, both Vandagraph, no cap required.	5 days
323732	15 Mar 2024	Derek Lamb	26 Mar 2024	0012176: Requires Cleaning! - Chipox fault, return to supplier; 0110452: Requires Cleaning! - Unstable and non-linear, normal sensor fault, no cap required.	13 days
323008	08 Mar 2024	Derek Lamb	26 Mar 2024	0110229: Requires Cleaning! - Low output; 0110452: Requires Cleaning! - Unstable and non-linear, both normal sensor type failures, 0110229 Vandagraph sensor.	20 days
320903	16 Feb 2024	Derek Lamb	23 Feb 2024	0111235: Oxygen monitor with alarms, labelled as an alarm fault, turned out to be a physical connector fault, no cap required.	9 days
317562	12 Jan 2024	Derek Lamb	19 Jan 2024	0111260: Requires Cleaning! - Display fault, being returned to supplier; 0310100: Requires Cleaning! - Keyboard fault, no cap required.	9 days
316896	05 Jan 2024	Derek Lamb	11 Jan 2024	0380030: Tom Thumb Infant Resuscitator, replaced gauge face, no cap required; 1410000: Foetal Heart Simulator model V1000, serviced, no cap; 8010006: Oxygen Sensor - rEvo CR22D, zero offset error, sensor fault, no cap.	7 days
315219	15 Dec 2023	Derek Lamb	21 Dec 2023	8010050: Requires Cleaning! - Electrolyte leakage, rare fault, will monitor for recurrence.	7 days
311418	03 Nov 2023	Derek Lamb	10 Nov 2023	0111250: Requires Cleaning! - Handi Oxygen Analyser, no cap required, returning to supplier.	8 days
310083	20 Oct 2023	Derek Lamb	20 Oct 2023	0111275: Oxygen analyser OxiQuant S, high reading, return to supplier, no cap required; 7910110: Requires Cleaning! - No reading, Vandagraph, no cap required.	2 days

### Summary of Actions:

### Summary:
- The task of cleaning up product failure codes is systematically progressing under Derek Lamb’s oversight.
- Most issues were investigated, tagged, and closed promptly, with clean codes applied as required.
- Several products required no further action as no faults were found (e.g., routine recalibration, sensor checks, or return to stock).
- All stock references are being reviewed for new faults, and actions are being taken to categorize them correctly.
- Outstanding issues involve minor manufacturing defects or service adjustments, with actions pending supplier responses or additional investigation.
- Average task completion time ranges from 5 to 15 days, depending on issue complexity.
This review confirms that the process is effective, and the list is nearing completion.
Reviewed and all happy.
We do not have our own products. We do not have any concerns with regards to our suppliers products at this time.

No other issues. All happy.
### Follow on Issues

### Conclusion
We are at 28.3% of turnover. Roughly on target, we need to increase prices to to bring this back in line. We need this current level to maintain supply to our customers. No concern.
All happy with current. Prices to be increased

# (13.1) Storage and Stock Control
Objective: Suitable levels of Stock and Storage space. Check no issues have been brought up.
### Report References
Ref #336857.14654 Warehouse temperatures

### Notes
Current Stock Temperature Limits: -20 °C to 50 °C
Temperature (as of 10 October 2024, 4:21 pm):
  - Location: Main Office Stock Room, Goods Out Back Wall
  - Temperature: 20.60 °C
  - Pressure: 99895.00 Pa
Concise Summary
The warehouse stock room's temperature, recorded at 20.60 °C on 10 October 2024, falls well within the acceptable storage range of -20 °C to 50 °C. Historical data from April to October 2024 shows stable temperature trends, with monthly minimum temperatures ranging from 14.00 °C to 17.10 °C, and maximum temperatures between 21.50 °C and 26.20 °C. These values indicate a consistent climate conducive to safe stock storage, with no extremes recorded that would threaten stock stability.
No issues with temperature, damp or pressure. No extremes of temperature that would affect stock.
There are no space issues, everywhere is clean and tidy.
No need for extra stock space at this time.
All happy with stock and storage.

### Follow on Issues

### Conclusion
All happy with stock, no further action required.

# (13.2) Calibration Index Review
Objective: Calibration Index up to date

### Report References
 Ref #336857.14655 Calibration Index
Notes
 CE0185 (TF930 Frequency Generator)
  1. Status: Still with Cathy, awaiting calibration or replacement.
  2. Issue Dates:
    1. Outstanding as of 03 Oct 2024 (latest update).
    2. This equipment has been mentioned in multiple issues, dating back to June 2024.
  3. Notes: Philip Crossley has reported delays in its return and is awaiting resolution. In December 2023, it was replaced by CE216 and planned to be kept 6 months out of sync with CE0185.
2. CE082 (DVM) and CE111 (DVM)
  1. Status: Due for testing.
  2. Issue Date: 03 Oct 2024.
  3. Notes: Both devices are highlighted for testing, with plans to be processed soon.
3. CE210 (Flow Meter) and CE215 (Pressure Gauge)
  1. Status: Both need to be sent away for calibration.
  2. Issue Date: 03 Oct 2024.
  3. Notes: The flow meter has been previously reported for calibration since October 2023.
4. Cable Testers: CE188, CE191, CE194, CE195, CE196
  1. Status: Need to be located and checked.
  2. Issue Date: 03 Oct 2024.
  3. Notes: Philip Crossley noted that these testers are pending verification and calibration.
5. CE214 (Torque Screwdriver)
  1. Status: About to run out of calibration, suggested to replace it as it is cheaper to buy a new one.
  2. Issue Date: 03 Sep 2024.
  3. Notes: Decision to wait until it's needed, as it could be months before its use.
6. CE178 (Digital Pressure Gauge)
  1. Status: Recently back from calibration (Jul 2024).
  2. Issue Date: Previously reported in July 2024.
  3. Notes: Calibration is completed; no further action required.
7. CE164 (Microcal 1030)
  1. Status: Taken out of service due to being faulty.
  2. Issue Date: 05 Aug 2024.
  3. Notes: Microcal returned from calibration but was faulty and decommissioned.
### Summary of Actions:
- Immediate Calibration Required:
  - TF930 (CE0185): Pending calibration since June 2024.
  - DVMs CE082 and CE111: Testing needed.
  - Flow Meter CE210 and Pressure Gauge CE215: Send for calibration.
  - Cable Testers CE188, CE191, CE194, CE195, and CE196: Locate and check.
- Equipment to Monitor:
  - CE214 (Torque Screwdriver): Replacement needed, waiting until required for use.

Reports reviewed and the calibration index is up to date.
We have received a non conformance on this and we will be updating the VOP and adding date stickers to all calibrated units.
2560663-202410-N1 claus 7.6 Viamed ISO re certification from 31st October 24.

All happy with the calibration Index and calibration equipment status.
### Follow on Issues

### Conclusion
Calibration Index up to date.

# (13.3) Process performance
Objective: Review Process performance, is everyone happy with this.
### Report References
Ref #336857.14664 Employee Processes
Ref #336857.14665 Process reviews
### Notes
Reports reviewed, no issues with the process listings.
All process risk assessments has been completed.
No overdue outstanding process.

Product performance. We no longer have our own products.
All process are monitored, no issues with the processes or system.
Matters arising were all checked ok and found to be effective.
No details missing form the Audit roles and titles page, reviewed page and all happy.
Reviewed old issues, all happy.
### Follow on Issues

### Conclusion
Reviewed process performance. All happy with current system.

# (14.0) Back Orders
Objective: Time to push customer order through building. Review current back orders.
### Report References
 Ref #336857.14668 Active List as at 14 Oct 2024
Ref #336857.14669 OverView
Back order and forward orders live in Intrastats.
### Notes
Reports and active list reviewed
We use live data not uploads.
Back Orders report £129027.68
Forward order Report £41276.43
Reviewed days to ship, no issues.
No issues with the Active List. It had a small update. Added method to ship documents with orders.
No Problems. All happy with current status.
### Follow on Issues
Issue 345838 and Issue 345839 sent check the back order amount, as it seems high. To verify its still working as intended.
### Conclusion
No issues other found, all happy with current status.

# (15.0) Customer Complaints
Objective: To review / close any customer complaints within 30 Days. Review all customer complaint from the last 12 months.
### Report References
 Paper file customer complaints - no complaints.
Notes for discussion:
Handling of Customer Complaints where we are ‘In the Middle’ I.e. we are a conduit to the supplier from the customer. Timescales may be beyond our control,
See issue 304005 customer commented on a bluepoint product, we rasied an internal complaint, but its not solvable by us, and indeed is user abuse of the connector,
bluepoint took ages to send us a letter saying this is the only occurance, so no changes have been made to the product.
Do we add a section to our procedures in dealing with being in the middle specifically?
Reviewed reports and issues.
304005 – issues resolved with Bluepoint
This took 12 months to close as was a supplier issues. How should we be dealing with this moving forward. So they do not go on so long.
We discussed the issues and it has been decided to add a sectioning in the VOP - complaint from customer regarding supplier product. We will not raise internal Viamed customer complaint. Will create issue in header suppler complaints us to them. To log all communications. Keep open until customer happy.
Nothing not already reviewed, all happy.
No other complaints in the last 13 months.

All reviewed and all happy with current statue of the above customer complaints.
### Follow on Issues
Added to existing issue 345404. To update the VOP and add supplier complainant to the complainants review. Also see issue 345844.
### Conclusion
All complaints reviewed from the last 13 months. All new customer complaints are reviewed or completed within 30 days. All reviewed and all happy with current status of the above customer complaint.

# (15.1) Customer and product feedback
Objective: Review Customer feedback over the last 12 month. Check none should have been a Complaint.
Review Product Feedback from customers

### Report References

 343734 (14 Oct 2024): Sales forecasts and warnings for specific products with significant sales drops, including oxygen sensors and a foetal heart simulator.
342626 (02 Oct 2024): Feedback from Steve Collard regarding a website order for a Divesoft O2/He analyser.
342607 (02 Oct 2024): Request for feedback from a unique customer regarding EyeMax Flowsensors.
342568 (01 Oct 2024): Feedback from Royal Preston regarding equipment servicing issues.
342138 (26 Sep 2024): Feedback on R-45V oxygen sensor performance.
340944 (16 Sep 2024): System-generated sales forecasts for sales warnings.
338790 (21 Aug 2024): Positive feedback from eBay regarding MD300-C29 Finger Oximeter.
338787 (21 Aug 2024): Positive feedback from eBay for MD300-C5 Frog Paediatric Finger Pulse Oximeter.
338730 (20 Aug 2024): Feedback on rainout issues in Hamilton ventilators.
338095 (14 Aug 2024): System-generated sales forecasts for stock warnings.
336244 (25 Jul 2024): Positive feedback regarding R-22VA oxygen sensors.
335640 (18 Jul 2024): Positive feedback from Nottingham University Hospitals on aftersales care and support for MaxO2 ME sensors.
333541 (26 Jun 2024): Feedback about unusual failure of R-43V oxygen sensors from Alder Hey hospital.
330051 (20 May 2024): Positive feedback from James Cook University Hospital on spare parts and stock availability for Tom Thumb devi

Summary:
October 2023
1. 19 Oct 2023:
  1. Customer Feedback: The customer accepted a 50% credit as a goodwill gesture (EUR 145.95) for an issue with Medisera products. The credit was applied, and the customer thanked Viamed for the resolution.
2. 18 Oct 2023:
  1. Product Feedback Review: A general review of customer feedback for the past 12 months found no complaints among the feedback gathered.
3. 16 Oct 2023:
  1. EyeMax Feedback: Feedback regarding EyeMax products, noting a satisfaction survey for EyeMax phototherapy masks. Customers were generally satisfied with the product.
### November 2023
1. 10 Nov 2023:
  1. Customer Feedback for Maxtec EyeMax: The feedback included mixed responses. A customer from the Panda unit and A&E pediatric department rated the product highly, though they mentioned the product as "overpriced" and the value for money as average.
### January 2024
1. 05 Jan 2024:
  1. EyeMax Feedback: Mid and South Essex NHS Foundation Trust provided highly positive feedback on EyeMax masks, noting reliability and high quality of the product.
### Additional Feedback:
- 02 Oct 2024: Positive feedback from Steve Collard on excellent service for a website order for a Divesoft O2/He analyzer.
- 02 Oct 2024: Royal Preston provided feedback regarding equipment servicing.
### Notes
The Board reviewed all the feedback, and we are all happy with the feedback. No problems have arisen from the feedback we have received.

We are now receiving feedback from Survey Monkey which is going in to the system. Now we are getting more responses from customers we can look at linking this feedback back to the supplier. To assists them with their Standards.

None of the feedback we have received is a complaint. There are no ongoing issue.

No other issues, all happy with feedback.

### Follow on Issues

### Conclusion
No issues found in customer feedback.

# (15.2) Non Conformities Review
Objective: To review and action on any non conformances within 15 days. Check for new qc 21 forms.
### Report References
Ref #336857.14671 Complaints and Non Conformance
Order Errors:
- Incorrect order entries were logged, such as address errors (Issue ID 343889) and carriage misprocessing (Issue ID 343820). These were identified as internal office processing mistakes.
- Another error involved shipping items incorrectly using UPS instead of Royal Mail as per customer memo instructions.
Shipped Item Returns:
- Multiple return cases were recorded, including a box return due to a manufacturing defect involving loose PCB components (Issue ID 343727).
- Similar return incidents were logged, requiring non-conformance reviews to determine if CAPA (Corrective and Preventive Action) processes should be initiated.
Non-conformance Reviews:
Regular reviews were conducted on non-conformance logs (e.g., Issue ID 343025), checking actions implemented to prevent recurrence and monitoring any increase in error rates. These were reviewed to ensure corrective actions are in place and effective.
Customer Complaints:
- Complaints were filed regarding system or staff errors, such as an issue with the Shopify Admin Page (Issue ID 341548), where an order was not fulfilled. It was resolved by internal review and corrective measures.
- High QA failure rates were noted in some reports, requiring attention and review by relevant staff (Issue ID 342609).
ISO and Training Audits:
Audits identified issues such as incorrect task numbers in training documents and were subsequently corrected (Issue ID 341465).

### Notes

Reviewed the above reports reports. No Issues.
Order errors reviewed and no issues all ok.
Reviewed all non conformance. New QC21 form has been generate as Expired document where not being tracked. This was picked up internally and a form generate. Action plan filled in. Everything else has been dealt with. No issues.
QC21 form – issues 345017 ongoing. Nothing else ongoing. All happy with the current reports and status.
All non conformance are reviewed and actioned within 15 days.
As discussed above BSI N1 has just come in and will be entered on to a QC21 form imminently. The was a calibration labeling issue.
### Follow on Issues
QC21 form current issue 345017

### Conclusion
Reviewed all non conformances, no issues or areas of concern. All happy with current status.

# (15.3) Vigilance System
Objective: Assess whether any reporting needs to be made that has not already been reported over the last 12 month. Ensure the board is happy with existing vigilance issue for the last 13 months.
### Report References

Ref #343435.14652 Product Fail Reviews- failure codes report live in Intrastats. Rolling task 750 Product Failures New Codes Review.

### Notes
Nothing to report all happy with status of vigilance issues over the last 13 months.

Product failure codes report reviewed.

### Follow on Issues

### Conclusion
No Issues with vigilance. No reporting needs to be made that has not already been reported over the last 13 months.
The Board is happy with the existing vigilance issue for the last 13 months.

# (15.4) Preventive and corrective actions
Objective: Review Risk Benefit Processes Report. Check any corrective action required. Are there any new risk identified.
### Report References

Global overall Risk Assessment Review Version 1730727678. Reviewed live report in Intrastats.
Effectiveness Review – Rolling Task and Audits Page. Reviewed live report in Intrastats.
### Notes
Reviewed Risk benefit Report in the Audit, Roles and Titles heading. No corrective action required. The Risk Benefit reports. Was reviewed and all is ok no issues.
The Process Updates Effectiveness Review was carried out in the meeting by Derek Lamb.
The Rolling Task and Audits Page was also reviewed in meeting. It is used to prevent adverse effects on the system.
See Report Review Risk Benefit Processes Report made so we can assess risk more easily. We are constantly maintaining and improving our systems. We have utilised feedback from staff, customer and suppliers to prevent errors and where errors are reported we carry out internal reviews to ensure they do not re occur. Our system is design to be improved, to fit in with new standards and requirements. Intrastats has check sums in place to prevent mistakes. These are also being added all the time, when we find areas that would benefit from them. Corrective actions come in the form of software developments, issues and training where needed.

Process updates – effectiveness is reviewed monthly. All ok no issues.
No other new risk identified with old products no action required. At this time no new Risks have been identified. All happy.
Old technical files and PMS are up to date. No problems in the last 12 months.
### Follow on Issues
### Conclusion
No problems have surface in last 12 months.

# (16.0) ISO Issues
Objective: Review Notifications, Recalls. Ensure all Regulatory requirements are up to date.
### Report References
### Ref #336857.14709 BSI 2391793-202310-N1
Servicing process was not fully effective as servicing procedure
and information was not seen to include analysis to determine if
information is to be handled as a complaint.
Clause 7.5.4
VOP09, 24. 11. 2021 Repairs and Servicing
SRN 36055, Tom Thumb SN 0401370, 11. 10. 2023
SRN 36045, barcode 987618, M000SN 4278, Microstim DB3, Report
1697639395, 18.10. 2023.
Added pull down for engineers to use to highlight potential Nonconformance or customer complaints from the returns service.
Checked it shows in the complaints review.
Viamed Scope reviewed.
** Update the above Non conformance has been closed by BSI on the visit 30th and 31st October 2024. **
Created new non conformance from the new assessment. 2560663-202410.N1. Just been received 4th November 24.
### Notes
More global for discussion, is we will be reducing our scope with BSI regarding discontinued products, we now out of service support for most of our old range, do we keep 13483? do we move to 9001?
Agreed to drop first sentence of scope service of Supra-Maximal nerve simulator, simulation equipment etc.

Just leave Distribution and Service of Medical Devices and Products. Derek will discuss with BSI scheme manager.
No Notifications and no recalls in last 12 months. All Regulatory requirements are up to date and reviewed regularly. All happy no issues. No supplier recalls. Viamed no longer has its own product.
### Follow on Issues

Issue sent 345870 reduce scope, Derek Lamb to contact scheme manager.

### Conclusion
Reviewed the Notifications and Recalls. All Regulatory requirements are up to date. No problems all happy.

# (16.1) Regulatory Changes
Objective: To review any standards we hold, for regulatory changes.
### Report References
Task 48
Reviewed and no changes to the 13485:2016.
 ISO 13485-2016 Still Current as per our certification.
### Notes
No notifications.
No Recalls.
Task 48 review Regulatory Requirements is up to date.
Reviewed and no changes to the 13485:2016.
ISO 13485:2016 Still Current as per our certification.

We are UKRP for Nufer, Sunnybaby, Ornatec and Teledyne. We are MHRA registered for these. This is ongoing but reviewed regularly.

UKCA has been delayed until 2028.

### Follow on Issues

### Conclusion
Reviewed, no issues.

# (16.2) Documentation and records
Objective: Check everyone is happy with current Document Index. Verify new Risk is being assessed in the document update process.
### Report References

Statistic Header	------------------
Barcodes	2,525,538
Barcode Movements	2,881,191
Documents	165,516
Live Documents	77,645
Access Logs / Page Impressions	1,401,514
Tracking Customer Orders	23,626
Total Issues	335,999

.
Tracking 12% more barcodes in 2024,
26% increase in document storage,
25% increase in number of live documents
Tracking Customer orders is broken as same report as last year,
new Order System with:
(32657 total, Query took 0.0003 seconds.) records,
The Overview page Needs a complete rewrite, as its reading old opera tables,
underlying data is 100% ok, its just the overview report screen

### Notes
The above figures were reviewed and no issues.
VOP 01 needs updating to remove on site servicing. Issue sent 345397 to Helen lamb.
No problems with the documents and records digital system.
All happy with Document index. Nothing is ever deleted and all show newest first. Everything is backed up off site.
Risk assessments are being carried out when Documents are being upload and updated. Also when any updates occur in the system.

### Follow on Issues
345397 Issue to Update VOP 01 to remove on site servicing.
### Conclusion
All happy with current state of document index, and risk reviews on updates.

# (16.3) Quality planning
Objective: Review Quality planning and ensure everyone is happy to continue with current Quality policy.
### Report References
Ref #336857.14690 VM3COP00.00 VOP00.00 Viamed Quality Statement policy and objectives
Ref #336857.14677 Responsibility's Organsiation Chart
Ref #336857.14666 ISO 13485 Scope for review
Ref #336857.14694 Objectives Expanded
Ref #336857.14667 Exclusions To 13485 Review

### Notes
Reports Reviewed
Viamed Quality Statement policy and objectives was reviewed and all happy.
Reviewed Responsibility's Organsiation Chart, needs Humanmed removing issue sent. All happy with the rest.
Viamed ISO 13485:2016 Scope was reviewed, updating this with scheme manager.
Objectives and Achievements was reviewed and all happy.
All happy with Boundaries / Exclusions, VM3COP 02.01 Document ID 74571.
All happy with the BSI official Certificate.
Agenda was reviewed and all happy with it. We are adding issues when we find objectives that need to be updated or added.

### Follow on Issues
Issue sent 345880 to Derek Lamb to remove Humanmed from the orgaisation chart and re print.

### Conclusion
Everyone is happy with Scope, Quality Policy, Exclusions / Boundaries. Changes to be made to the official BSI Certificate. Issue being sent where objectives need updating. Everyone was happy to continue with the current Quality policy and Quality planning.

# (16.4) Achievement of Quality Policy and Objectives.
Objective: Achievement of Quality Policy and Objectives. Confirm still suitable and happy to continue with current.
### Report References
Report from BSI is pending. No report reference.
Objectives and Achievements reviewed.
### Notes
As of 31st October external BSI audit has recommended a continuation of our certification.
Therefore we have achieved our quality policy and objective.
Objectives and Achievements reviewed and all happy.
All agreed happy to continue with current Quality policy.
We have renewed our certification, so we have achieved the objective. All agree the Quality Policy and Objectives are still suitable.

### Follow on Issues

### Conclusion
We have confirmed the objectives and policy's are being achieved. All confirmed this is still suitable and we are all happy to continue with the current Quality policy and Objectives.

# (16.5) Changes to the Management system
Objective: Review any Changes to the Management system. Ensure everyone is happy with current system. Assess the effectiveness of the system.
### Report References
 Ref #336857.14707 ISO 13485 2016 Route Map

### Notes
Report Reviewed
ISO 13485 Route Map reviewed.

Effectiveness of the system has been assessed.
We are documenting risk analysis throughout the system. We are carrying out Risk assessments to prove changes do not affect current ISO system or require external notification.

There have not been any changes to the Management System.
Everyone is happy with the current Management system.
The Board finds the Management System to be effective.
All happy with the effectiveness of system

### Follow on Issues

### Conclusion
All happy with the effectiveness of the system.

# (16.6) Changes to CE marked products
Objective: Review any Changes to CE marked products, current or those in their lifespan.
### Report References
Viamed CE Marked Medical products have all been discontinued.
No changes to legacy products. Or CE products in their lifespan.

### Notes
There are now no products in their lifespan.
 Viamed CE Marked Medical products have all been discontinued.
### Follow on Issues

### Conclusion
No Changes and no Viamed CE Marked Medical products. Reviewed and All happy with current.

# (16.7) External Parties risks and opportunities review
Objective: Review External Parties risks and opportunities. Confirm all happy with current.
### Report References
Ref #336857.14706 External Parties
Need to action the new supplier review questionaire, also need to ensure the question of climate change is addressed, its been added to ISO 9001:2015 as an amendment, so can foresee it being added to 13485:2016.
Added the clause to the review, still to be addressed.

### Notes
Report Reviewed.
Reviewed external parties Risk and Opportunities.
Reviewed the list of external parties all happy with current.

### Follow on Issues

### Conclusion
Happy with current report.

# (16.8) New products
Objective: Discuss any possible new products.
### Report References

### Notes
No Viamed products. It is based on suppliers releasing new products.
We are not designing our own products anymore.

### Follow on Issues

### Conclusion
No new products.

# (16.9) Advisory notices & recalls
Objective: Review any Advisory notices or recalls of our products.
### Report References

Reviewed Intrastats Meeting Header Advisory notices and recalls
### Notes
No advisories, notices or Recalls.

### Follow on Issues

### Conclusion
No advisories, notices or Recalls.

# (16.10) Possible system changes
Objective: Review any system changes.
### Report References
Objective
To outline potential system changes and provide an update on current server status.
Report References
A new main server and backup server have been installed.
Server Overview:
- vmserver10.thevault.me.uk – Live working server.
- vmserver9.thevault.me.uk – Live backup server (offsite, mirroring live system).
- vmserver10.viamed.co.uk – Live development server, ready to act as a swap-out for both hardware and data when needed.
Amendment log Reviewed.
### Notes
No technical changes, we have been updating our hardware, new main server and backup server, both now in place.

Staff PC’s are in process of being updated. Computers bought in 2020 for remote working for Covid are now wearing out. So we are replacing these.
Nothing major, continual improvements, always on going.
No major system changes planned.
Planning a second internet line in to the building, to maintain access to the systems.
Discussed the amendment log. No areas of concern.

### Follow on Issues
### Conclusion
Intrastats under continual gradual improvements. No major system Changes Planned.
# (16.11) Recommendations or improvement
Objective: Review and discuss any Recommendations or improvement
### Report References

### Notes
Recommendations or improvement were discussed and it was confirmed that none are required at this time.
Ongoing process in relation to Intrastats.
All happy with current.

### Follow on Issues

### Conclusion
No fundamental changes required.

# (16.12) Results of internal audits / Mini Audits
Objective: Review the results of internal audits / Mini Audits over the last 12 months. Check everyone is happy with the status of the current years Audits. Reviewed Audit scheduled for next year.
### Report References

Ref #343435.14674 audit results chart
Ref #343435.14675 Results Mini Audits
Ref #343435.14673 Audit cross reference Audit calendar
Ref #343435.14770 QC17 Outstanding Audit and current status
### Notes
Audit scheduled, no fundamental changes. Just updated with the new processes when carried out for the current year
Reviewed mini Audits over the last 12 months. No problems all happy.
Reviewed internal Audits over the last 12 months, all happy.
Reviewed summary of 2023 audit next actions, all happy.
Reviewed Outstanding Audits and current status, all happy
All running a smoothly and up to date.
All agreed happy with current years Audits and all happy with the Audit schedule for 2025.
### Follow on Issues

### Conclusion
Happy with 2025 Audits of the system.
Happy with current status of 2024.
All Agree to the 2025 Audit schedule as per VOP/Audit QC17 Audit calendar in the route map page.

# (16.13) Audits Meeting Closure
Objective: To agree the current schedule of Top Level audits

### Report References
 Ref #336857.14673 Audit Cross Ref Check and Calendar
### Notes
Report Reviewed.
Audits Calendar Viamed was reviewed, no problems.
All agree to current Schedule of top level audits.
All agreed with 2023 Audits results. Closed 2023 Audits.

### Follow on Issues

### Conclusion
All Agree to the 2025 Audit schedule as per VOP/Audit QC17 Audit calendar in the route map page.

# (17.0) Company Issues
Objective: Review Company Issues.
### Report References

### Notes
No company issues.

### Follow on Issues

### Conclusion
No company issues. No areas of concern raised.

# (17.1) GDPR
Objective: Review GDPR and ensure it is current.
### Report References
Ref #302478.13292 Viamed Group of Companies GDPR Privacy Policy ID113160

### Notes
Report Reviewed
Viamed Group of Companies GDPR Privacy Policy ID113160 no issues.

There have been no changes or updates, this was reviewed and all happy with policy and current status.

### Follow on Issues

### Conclusion
Reviewed and happy with that the current policy is up to date.

# (17.2) Review of responsibilities
Objective: Review the organisation chart. Check everyone is happy with current responsibilities.
### Report References
Ref #336857.14677 Responsibility's Organisation Chart
Ref #336857.14676 Responsibility's review
### Notes
Reports Reviewed
Reviewed the Viamed Company Responsibility's organisation chart structure everyone is happy with current. All agreed current organisation chart.
Reviewed Roles and Responsibilities, all happy with current. No changes needed. Looked at the individuals Roles and Responsibilities in Intrastats. All happy with current responsibilities.

All agreed happy with current.
### Follow on Issues

### Conclusion
Happy with current responsibilities.

# (17.3) Resources required
Objective: Do we need to change any resources, discuss all areas do they need any further resources. Review Equipment, Training, Staffing, Standards, Systems, Building fabric.
### Report References

### Notes
Review resources
Equipment – on going computer equipment improvements and repairs/updates. Fixing where needed. Some of the staff PC 's bought in 2020 are failing and will be replaced as needed. New laptops have been bought for the directors, further laptops will be purchased. Looking at a new cable stripping machine which has been agreed to buy one if appropriate. New Internet line for the company, to act as backup. New internet for the MD/IT director to ensure he is online at all times.

Training – ongoing, office training, procedures for new staff, product training. No current requirements.
Staffing – current levels ok at present. Possibly a little over staffed.
Standards – UKCA under review.
Systems – Nothing new required.
Building and fabric – adequate at present, noting new needed.

### Follow on Issues

### Conclusion
Reviewed and ok with the resources we currently have

# (17.4) Net Zero Carbon reduction
Objective: Review current Plan and targets, report on current status.

### Report References

### Notes
We are working towards Net Zero in accordance with our carbon reduction policy.
Paperless working. Off setting with help form the Woodland trust. Regular reviews.

### Follow on Issues

### Conclusion
We are starting out and working towards Net Zero. More will be done over the next 12 months.
# (18.0) Building fabric Issues
Objective Review Building fabric Issues.
### Report References

### Notes
The garden is in the process of being tidied up.

### Follow on Issues

### Conclusion
No other issues at this time. All happy with current building and fabric issues.

# (19.0) Responsibilities, Staff , Training Issues
Objective: check organisation chart is up to date. Discuss any possible training requirements. Discuss the effectiveness of training is it effective, are the staff competent. Is any further training planned.
### Report References

### Notes
Competence Office Staff and goods out Report reviewed for last 12 only minor issues, all happy. Staff Internal Training 2024 reviewed, no issues, all happy
The organisation chat has already been reviewed and found to be ok.
The training Requirements were discussed in resources.
The Effectiveness of training was discussed it was agreed it was effective for our current requirements.
Staff are competent in the roles they have been allocated. Effectiveness reviewed in the weekly review of rolling tasks and responsibilities.
All agreed training is effective and staff are competent.

### Follow on Issues

### Conclusion
Happy with Responsibilities, Staff and Training.

# (19.1) Staff Appraisals
Objective: All staff appraisals to be carried out within 18 Months of Last appraisal or first Hiring
### Report References
 Ref #343435.14771 Staff Appraisals
### Notes
Report Reviewed.
Staff Appraisals, completed and reviewed. No problems raised.
These were carried out during June / July 2024. All are now complete and have been reviewed and any issues addressed.
The appraisals where based on statistics to show any areas of concern. Then the effectiveness was reviewed. All staff were found to be compitent.

### Follow on Issues

### Conclusion
Appraisals have been carried out this year and all items raised were resolved. All happy.

# (20.0) Distributor Issues
Objective: Review Distributor Issues.
### Report References

### Notes
Oxylink and SUCH are very demanding customers. But at present they are worth it.

Carbon reduction Net zero was discussed regarding distributors and it was determined our time should be spent on our suppliers and not distributors at this time.
Trionara is overdue but we are chasing.
### Follow on Issues

### Conclusion
Happy with current status of distributors.

# (21.0) Supplier Issues and review
Objective: supplier reviews up to date
### Report References
Ref #343435.14650 Supplier Review Index

### Notes
All up to date. This is system forced, you cannot put an order on if the Supplier Review is out of date.
All happy with current system and suppliers.
We are in the process of re writing this process and system as we are required to have a lot more upstream monitoring of our suppliers with regards to human slavery, carbon net zero and other environmental monitoring. We anticipate this being in place in the next 6 moths.
A lot of suppliers are being sold and managements changing, so we are having to get to know them all over again and then we get issues. Sometimes they discontinue products and have issues shipping. We monitoring these changes.
### Follow on Issues

### Conclusion
All happy with current system and suppliers. Supplier review is up to date.

# (21.1) Supplier / subcontractor performance
Objective: Review any Supplier / subcontractor performance.
### Report References

### Notes
Reviewed above. We do not have any subcontractors.

### Follow on Issues

### Conclusion
No subcontractors.

# (22.0) Any other Business
Objective: Make sure everything is covered and renewed, ask all present to consider if any thing has been missed
### Report References
Ref #336857.14711 VM3COP60.04 Viamed Disaster Planning Business Continuity Plan Reports ID72870
 Ref #336857.14712 VM3COP60.06 BUSINESS CONTINUITY PLAN Contingency Plan Viamed VST Vandagraph ID74819

### Notes
VM3COP60.06 Business Continuity Plan needs updating to reflect the new systems.
VM3COP.60.04 Viamed Disaster Planning needs updating to reflect the new systems.

Reviewed health and safety areas, risk assessments, fire assessments and hazards.

We have checked Objectives, Effectiveness, Audit calendar and signed off last years audits. We have carried out the Appraisals and found them to be effective. We have assessed Risk when updating or changing the documents and processes. As a company we have become more aware of risk and assessing it
We have reviewed feedback, complaints handling, reporting to regulatory authorities, monitoring and measurement of processes, monitoring and measurement of product, corrective and preventive action, follow-up actions from previous management reviews, changes that could affect the quality management system, recommendations for improvement and applicable new or revised regulatory requirements.
We would like to monitor the effectiveness of Medica and the other exhibition.
We are in progress of getting more products, that can be shipped in and shipped out with ease.

### Follow on Issues
Issue 345953 sent to Helen Lamb to update the Disaster planning document
Issue 345954 sent to Helen Lamb to update the Business continuity plan

Conclusion
No problems or any other business.

# (22.1) Review Management Meeting Headers
Objective: To make sure no areas have been missed. Is the agenda suitable and up to date

### Report References

### Notes
We have added issue 345955 to add a Health and Safety heading in to the Management review. With the objective – To confirms risk assessment and health and safety assessments are still valid. Documents to be added to header - escape routes, fire risk assessment, health and safety assessments.
Agenda Headers have been agreed – these were assessed by the Board and found to be complete and effective. We feel that they cover all areas of the company and its systems. No other changes proposed. We are happy that the objectives are fit for purpose.
### Follow on Issues
Issues sent 345955 to Derek Lamb to add a new header and objective regarding health and safety.

### Conclusion
Happy with current headers and objectives, and feel all section of the company have been covered. Once the new Health and safety header has been added.