VM3COP50.11 Fitting a resuscitation bed 0320019 / 0320021 


1.
Open the resuscitation cabinet and lower the door to the horizontal position. Measure 5cm from the left-hand edge of the door at the back near to the hinge plate and mark the position.

[image: image4.jpg]



2. 
Measure 5cm from left-hand edge towards the front of the door and mark the position.

[image: image2.jpg]



[image: image1.jpg]e

—




3.
Align the left-hand side of the bed with the marks and align the front lipped edge of the bed with the front edge of the door, ensuring that the front edge of the bed does not protrude beyond the edge of the door, as this may cause the bed to foul on the cabinet body when closing.

4. 
Using a drill with a pilot drill bit, drill a small pilot hole approximately 5mm deep in the centre of the mounting 6 holes in the bed. Fix the bed into place using M6 x 1/2” pozidrive pan-head screws, being careful not to over-tighten so as to avoid cracking the bed.

[image: image3.jpg]



12/08/13   

Page 1/2

