QC 21 Non Conformance Report

	Date
	11 September 2025

	Viamed / VST* Issue id
unique identifier 
*delete as appropriate
	375175

	BSI Ref (if applicable) 
unique identifier 
	2700663-202509-N2

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The internal audit process is not fully effective as raised issues within audit records do not clearly differentiate between observations and non-conformities.

Evidence:

 Audit records reviewed (Audit 15 and Audit 06) highlighted issues and were raised within the Intrastat (eQMS system). The issues raised did not highlight or state if the issues were an observation or non-conformance. No investigation or CAPA related actions were seen and the issues were cleared/completed with note comment only. # Audit 06 – Performed 30 May 2025 – Issues 365729 Raised 26 June 2025; Issue 367474 Raised 18 June 2025 – Closed 26 June 2025 # Audit 15 – Performed 21 July 2025 – Issues 370409, 370410 – Raised 21 July 2025 – Closed 22 July 2025

	Investigation By:
Person responsible 
	Derek Lamb, Helen Lamb

	Investigation
Root Cause Analysis 
Issue id (if applicable)

	That we weren’t distinguishing between observation issues and non conformances when adding follow on issues to audits. 

When completing the follow on issues from the audits there wasn’t anywhere near enough details added to the issue to say what had been done and the problem. 



	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action 
Issue ID (if applicable):
Relevant Immediate Corrective action (if applicable) 
	Immediately start adding to follow on issues, if the issue is an observation issues or a non conformances issue on Audits that are due or coming due now. 



	Time Scale for Immediate Corrective Action
Time for completion of all identified actions 
	1 Week - 18th September 2025

	Corrective Action:
Relevant and Proportionate Corrective Action
	To add, to the Audit issues that are sent out monthly, for all the audits, a statement that says “Any follow on issues must be identified with Observation Issue or a Non Conformance Issue.

If the issue is a Non Conformance a QC21 form must be filled in, as per the QC21 form procedure, and investigated as such”

We will update the VOP 13 to reflect this. 

Send out an issue to all staff to advise them that when completing issue that relate to non standard (recurring) issues to add more information on completion. To say that a task has been completed satisfactory and what was done where relevant. Also if it requires a follow ups or if its a non conformance. 



	Time Scale for Corrective Action
Time for completion of all identified actions 
	1 Month - 11th October 2025

	Risk Review 
Confirm action does not have negative affects and/or risks to any other part of the ISO system 
	No risk has been identified in relation to adding to the audit issues and the VOP. As this is an addition and expected to improve the task. 

	Follow-up future issue id

(Effectiveness verification)
	#375181 

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


