QC 21 Non Conformance Report

	Date
	11 September 2025

	Viamed / VST* Issue id
unique identifier 
*delete as appropriate
	375175

	BSI Ref (if applicable) 
unique identifier 
	2700663-202509-N2

	Responsibility
Person Overall responsible 
	Derek Lamb

	Non-Conformance
statement of the problem 
	The medical devices files process of the organisation is not fully effective, as the organisation has not established requirements to ensure that product labelling is up to date and viable

Evidence:

The organisation has not documented labelling requirements as a distributor to verify or maintain product labels implemented and placed by the legal manufacturers are current, compliant and remain legible throughout distribution. Technical files were seen historic, predating the sampled product and existing labels on product was seen not recorded. The product sampled was in relation to Temperature Probe – Skin Contact – Ref 1010132021V – SN 210114478 – Type 0212921 – Viamed Ltd – Produced Oct 2021 – Barcode ID 1856344 – Manufacturer Bluepoint Medical Gmbh # Interface Agreement – Bluepoint medical GmbH & Co KG – 21/12/2016



	Investigation By:
Person responsible 
	Derek / Helen Lamb

	Investigation
Root Cause Analysis 
Issue id (if applicable)

	We have missed the need to review stock that we have bought in from our suppliers, that hasn’t sold straight away (within a year). It is a process we have not done and not considered. 



	Corrective Action By:
Person responsible 
	Derek Lamb

	Corrective Action 
Issue ID (if applicable):
Relevant Immediate Corrective action (if applicable) 
	Put a hold on any stock that has been with us for longer than ***12 Months???****. So nothing can be sold that hasnt’t been verified as matching to the suppliers current labelling, IFU, Declaration of conformity or CE certificate. 



	Time Scale for Immediate Corrective Action
Time for completion of all identified actions 
	1 week, 18th September 25

	Corrective Action:
Relevant and Proportionate Corrective Action
	To collate a list of all stock that we have had in for longer than **12 months ??**. Then to contact all the relevant suppliers to verify their current labelling, IFU, Declaration of conformity or CE certificate matches what we have on and in our stock. 

To ensure we record all communications to allow traceability of this verification. 

If our stock is out of date then we will request our suppliers take back the stock to re label/update or ask they give us written permission that we can do the corrections in house. 

All these communications will be added to the CRM in Intrastats, so this verification can be easily retrieved. 

If any supplier does not reply to us regarding out of date stock we will place these stock in a holding location pending disposal. 



	Time Scale for Corrective Action
Time for completion of all identified actions 
	2 months, 11th November 

	Risk Review 
Confirm action does not have negative affects and/or risks to any other part of the ISO system 
	The risks would come from a lack of communications between supplier and us, and us and our staff. We will endeavour to record all communications and make sure when passing jobs to staff that these are all in the Intrastats system and on rolling tasks. So no miscommunications can occur and so we have records of the jobs being completed. 

With this we should minimise the risks of stock being missed, the wrong labels, IFU, Declaration of conformity or CE certificate being used and therefore the risk of goods being shipped with the wrong details on or inside should eliminated. 


	Follow-up future issue id

(Effectiveness verification)
	

	Effectiveness verification
	

	Closed By:
	

	Closed on
	


