QC 21 Non Conformance Report

Date	11 September 2025
Viamed / VST* Issue id / unique identifier / *delete as appropriate	375175
BSI Ref (if applicable) / unique identifier	2700663-202509-N2
Responsibility / Person Overall responsible	Derek Lamb
Non-Conformance / statement of the problem	The medical devices files process of the organisation is not fully effective, as the organisation has not established requirements to ensure that product labelling is up to date and viable Evidence: / The organisation has not documented labelling requirements as a distributor to verify or maintain product labels implemented and placed by the legal manufacturers are current, compliant and remain legible throughout distribution. Technical files were seen historic, predating the sampled product and existing labels on product was seen not recorded. The product sampled was in relation to Temperature Probe – Skin Contact – Ref 1010132021V – SN 210114478 – Type 0212921 – Viamed Ltd – Produced Oct 2021 – Barcode ID 1856344 – Manufacturer Bluepoint Medical Gmbh # Interface Agreement – Bluepoint medical GmbH & Co KG – 21/12/2016
Investigation By: / Person responsible	Derek / Helen Lamb
Investigation / Root Cause Analysis / Issue id (if applicable)	Our QMS previously controlled labels when we acted as manufacturer/OBL. With the shift to distribution, procedures were updated for supplier certificates, DoCs and IFUs, but supplier labels were not included as controlled documents. This created a gap against ISO 13485 requirements for: / 7.4.3 (Verification of purchased product), / 7.5.3 (Identification), / 7.5.11 (Preservation of product), / 4.2.4 (Control of records). / Because labels were outside the controlled-document scope, there was no trigger at goods-in or during stock review to confirm the label version on file matched stock, or to retain an archived copy.
Corrective Action By: / Person responsible	Derek Lamb
Corrective Action / Issue ID (if applicable): / Relevant Immediate Corrective action (if applicable)	Put a temporary hold on any stock that has been with us for longer than 18 Months. So nothing can be sold that hasn't been verified as matching to the suppliers current labelling, IFU, Declaration of conformity or CE certificate.
Time Scale for Immediate Corrective Action / Time for completion of all identified actions	1 week, 18th September 25
Corrective Action: / Relevant and Proportionate Corrective Action	Add supplier labels to the controlled-document index (ISO 13485 4.2.4). these will be stored as scanned/PDF copies. Old superseded versions will be archived. / Update the Goods-In procedure to show the current approved label on-screen when booking product in. / If the instructions for use are available to Goods In without opening the product, these will also be verified to the system. / Staff will complete a tick-box to confirm: / The label on the delivered goods matches the approved version (ISO 13485 7.4.3, 7.5.3). / The label is still legible and intact (ISO 13485 7.5.11). / Configure the system to flag any label revisions received from suppliers so the master file and Goods-In display stay current. / Train warehouse and QA staff on the new step and retain training records. / After implementation, sample stock at 3–6 months to verify that label copies exist, match stock, and confirmations are recorded. / Update VOP 01 Documentation & Records Control to include supplier labels as a document type, with an automated task to flag any on-hand stock whenever a label revision is added. This ensures labels are held, tracked and acted upon in the same way as IFUs, DoCs and certificates.
Time Scale for Corrective Action / Time for completion of all identified actions	2 months, 11th November
Risk Review Confirm action does not have negative affects and/or risks to any other part of the ISO system	Integrating supplier labels into the controlled-document index and adding a verification step at goods-in reduces the risk of obsolete or illegible labels being present on sellable products. Potential residual risks are: / Supplier communication delay when a label revision is issued. / Human error if staff do not complete the tick-box verification. / IT failure affecting the label display function on the goods-in screen. / These risks are mitigated by: / Automatic flagging of new label versions against existing stock in the Document Index. / Mandatory tick-box confirmation at goods-in, supported by staff training, audit and updated procedures. / Periodic sampling of stock (>18 months) to confirm the label copy in the system matches physical stock. / Standard backup and recovery procedures already defined in VOP 01. / After 3–6 months, a sample review of transactions and stock will verify that label controls are effective and that no obsolete labels have been released.
Follow-up future issue id / (Effectiveness verification)
Effectiveness verification
Closed By:
Closed on