## Specific Goods or Services
Preliminary Market Engagement Supplier Questionnaire for Oxygen Therapy Products and Anaesthetic Breathing Systems
Framework Number | 6012040
Framework Title | Oxygen Therapy Products and Anaesthetic Breathing Systems
Notes and Instructions
This Questionnaire will be used to assist HSCNI Trusts in gaining a better understanding of the supply base to comply with the Specification required
Supplier Contact Information
Supplier Name: | Viamed Ltd.
Name of Contact Person: | Catrin Hollings
Telephone Number: | 1535634542
Contact Email Address: | catrin.hollings@viamed.co.uk
Lot Number and Title | Product Specification | Confirm adherence to the requirement by a Yes response or provide comment explaining why this requirement cannot be met.
Lot 1 - Oxygen Connecting Tubing | Must have two narrow bore connectors with ID of 4mm
Must be made of disposable medical grade PVC
Must be between 180cm - 210cm in length
Tubing must be anti-kink
Must be individually packaged
Oxygen Connecting Tubing - Narrow / Narrow. Must have two narrow bore connectors internal diameter of 4mm
Oxygen Connecting Tubing - Narrow Wide. Must have one wide and one narrow bore connector with internal diameter 4mm
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 2 - Aerosol Tubing Smooth Bore | Smooth Bore
22mm internal diameter
Must be between 180cm - 210cm in length approximately
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 3 - Aerosol Tubing Corrugated | Corrugated
22mm internal diameter
Length 30M approximately
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 4 - Aerosol Oxygen Bubble Tubing | Bubble Tubing
Must have a minimum of 3mm internal diameter
Length 30M approximately
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 5 - Venturi Masks | Single patient use
Individually packaged
Must include colour coded valves for the following oxygen concentrations; 24% 28% 35% 40% 60% Oxygen Concentration to be clearly printed on each valve. Valves must be available individually or as part of a kit containing the full range listed above
Concentration to be printed on valve
Individual valves must be available or in a kit with above range included
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 6 - Facemask Oxygen Non Rebreathing | Reservoir Bag
Must include an anti-suffocation vent that must not be covered during use
Tubing Length: Approximately 2 metres
Available in both adult and paediatric sizes
Individually Packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 7 - Tracheostomy Masks | Individually packaged
Available in both adult and paediatric sizes
Must have standard 22m connector and adaptor compatitble with oxygen tubing
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 8 - Aerosol Facemasks | Must be transparent
Standard 22mm connector
Available in both adult and paediatric sizes
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 9 - Low/ Medium Concentration Oxygen Masks | Single Patient Use
Available in both adult and paediatric sizes
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 10 - Capnography Sample Lines | Must be compatible with Microstream CO2 System | YES
| Must be able to connect to standard ventilator tubing | YES
| For short term use | YES
| Long term use for humidified environment | YES
| Sampling lines should be available in a range of lengths, typically up to 4 metres, suitable for both short term and long term use in humidified environements | Standard Length available 2.0m
| Individually packed | YES
| Conformity Assessment Markings relevant for N.I | YES
| Free from Natural Rubber Latex | YES
| Data Sheet Sheets/Technical Information to be provided | YES
| Instructions for Use to be provided | YES
Lot 11 - Nasal Cannula | Straight and or Curved prongs | Curved
| Must be 180cm - 210cm in length approximately | Standard Length available 2.0m
| Available in adult, Paediatric and Neonatal sizes | Adult and Paediactric only
| Individually packaged | YES
| Conformity Assessment Markings relevant for N.I | YES
| Free from Natural Rubber Latex | YES
| Data Sheet Sheets/Technical Information to be provided | YES
| Instructions for Use to be provided | YES
Lot 12 - Humidifier (Bubble Through) | Humidifier plastic component for set up
Safety Release Valve
Outlet Connection for oxygen tubing
Visible minimum and maximum water fill levels
Approx. vol 500cc
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 13 - Adult/ Paediatric Nebuliser (with Tubing) | Minimum capacity 5ml
Includes Tubing
Single use aerosol mask adult
Single use aerosol mask paediatric
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 14 - Adult Nebuliser (without tubing) | Minimum capacity 5ml
Without tubing
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 15 - Adult Nebuliser (Handheld) | Minimum capacity 5ml
With tubing
Hand held
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 16 - Catheter Mount | Connection elbow 22M/15F
Approximate length 170mm excluding connectors
Must have ported and non-ported options
Flexible
Components must be bonded
Individually wrapped
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 17 - Breathing Circuits | Must be available in a variety of lengths
The patient end (Y Piece) must be 22mmM / 15mmF connection.
The anaesthetic machine end must be 22mm connections.
The tubing must be of a suitable design / construction to limit collapse
The breathing circuit must come with a safety cap to protect the Y piece when not in use
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 18 - Mapleson C Circuit Adult, Single Patient Use | 6-9mm Spigot
Fresh Gas Connection
Fresh Gas Flow Tubing length approximately 180cm - 200cm
2Ltr Reservoir Bag
0cmH20 to 60cmH20 APL Valve
Angle Piece
22mm M / 15mm Patient Connection
Safety Cap
Single Use
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 19 - Mapleson C Circuit Paediatric with Gas Connection | 6-9mm Spigot
Fresh Gas Connection
Fresh Gas Flow Tubing length approximately 180cm - 200cm
0.5 Litre and 1 litre Reservoir Bag
0cmH20 to 40cmH20 APL Valve
Angle Piece
22mm M / 15mm Patient Connection
Safety Cap
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 20 - Mapleson C Circuit Paediatric without Gas Connection | Must connect to standard tapered connecter for a oxygen wall flow meter
Fresh Gas Flow Tubing length approximately 180cm - 200cm
0.5 Litre and 1 litre Reservoir Bag
0cmH20 to 60cmH20 APL Valve
Angle Piece
22mm M / 15mm Patient Connection
Safety Cap
Single Patient Use Symbol
Must be supplied with standard paediatric face mask
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 21 - Geudel Airways | Must be available in sizes 000 to 5
Must be colour coded
Must be individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 22 - Anatomical Inflatable Facemasks Adult | Must provide a range of adult sizes
Non Sterile
Single Use
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 23 - Anatomical Inflatable Facemasks Paediatric | For babies under 2kg
For babies over 2kg
Must provide scented and non scented options
Infant size
Child size
Non sterile
Individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 24 - Nasopharyngeal Airways | Must have a safety flared end
Single Use
Individually wrapped
Available in sizes 3mm to 9mm Inner Diameter
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 25 - Breathing Circuits for High Flow Oxygen Therapy | High flow oxygen therapy circuit
Must be supplied in a kit with tubing and water chamber
Must include all connectors for the circuit
Individually Wrapped
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 26 - Nasal Cannulas/ patient connectors for High Flow Oxygen Therapy | High flow Oxygen therapy nasal cannula with 22mm standard tubing connector
Must have Head Strap
Must supply a range of sizes
High flow nasal cannulas must support flows up to 70L for adults and 25L for paediatrics
Individually Wrapped
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 27 - Nebuliser Kit | Nebuliser Kit must be compatible with standard 22mm adult ventilator tubing
Must be supplied in a kit with tubing adapter and vibrating mesh Nebuliser
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 28 - Anaesthetic Disposable Soda Lime Cannister | Must remove carbon dioxide from exhaled air in rebreathing systems
Must be compatible with at least one of the following anaesthetic machines listed below: Drager Perseus Maquet Flow-I GE carestation Avance CF Draeger Fabius Avance CS2 Please detail the machine your product is compatible with.
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 29 - Flow Sensors | Flowsensor to measure patient ventilator minute volume
Must be compatible with at least one of the following ventilator machines listed below: Draeger V800 Avance CF Avance CS2 Please detail the machine your product is compatible with.
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 30 - Water Trap | To remove waste water and or moisture from Gas sample line
Must have ability to measure patient gas
Must be compatible with at least one of the range of anaesthetic machines listed below: Draeger Perseus Maquet Flow-I GE Care Station Primus MRI Fabius Philips Expression MR400 Avance CS2 Please detail the machine your product is compatible with.
Must be Latex free
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 31 - Adult Bacterial Viral HME Filter for Theatre Use | Must Have Gas sampling port
Must have smooth edges and sealed housing
Filter Efficiency to be Greater than 99.9%
Co2 port Luer lock with cap with tethered cap
The Bacterial/Viral Filter H.M.E. must have a Max time of 24 hours
The Bacterial/Viral Filter H.M.E. must be supplied individually packaged as clinically clean
The Bacterial/Viral Filter size does not impact on performance or functionality of Anaesthetic Machine/Ventilator breathing circuit
The connection of Sampling line to Bacterial Viral Filter must be secure
The connection of the sampling line to the filter results in minimal kinking of the line and has no impact on patient end
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the Anaesthetic machine/ventilator or breathing circuit
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 32 - Paediatric Bacterial Viral HME Filter for Theatre Use | Must Have Gas sampling port
Must have smooth edges and sealed housing
The construction of the Bacterial/Viral Filter H.M.E. must be transparent or translucent, non-conductive filter housing
Filter Efficiency to be Greater than 99.9%
Co2 port Luer lock with cap with tethered cap
The Bacterial Viral Filter H.M.E. must have a Max time of 24 hours
The Bacterial Viral Filter H.M.E. must be supplied individually packaged as clinically clean
The Bacterial Viral Filter size must not impact on performance or functionality of Anaesthetic Machine/Ventilator breathing circuit
The connection of Sampling line to Bacterial Viral Filter must be secure
The connection of the sampling line to the filter results in minimal kinking of the line and has no impact on patient end
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the Anaesthetic machine/ventilator or breathing circuit
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 33 - Neonatal Bacterial Viral HME Filter for Theatre Use | Must Have Gas sampling port
Must have smooth edges and sealed housing
The construction of the Bacterial/Viral Filter H.M.E. must be transparent or translucent, non-conductive filter housing
Filter Efficiency to be Greater than 99.9%
Co2 port Luer lock with cap with tethered cap
The Bacterial Viral Filter H.M.E. must have a Max time of 24 hours
The Bacterial Viral Filter H.M.E. must be supplied individually packaged as clinically clean
The Bacterial Viral Filter size must not impact on performance or functionality of Anaesthetic Machine/Ventilator breathing circuit
The connection of Sampling line to Bacterial Viral Filter must be secure
The connection of the sampling line to the filter results in minimal kinking of the line and has no impact on patient end
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the Anaesthetic machine/ventilator or breathing circuit
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 34 - Heat and Moisture Exchanger (HME) for long term ventilated patient | Must have smooth edges and sealed housing
The H.M.E. must have a Max time of 24 hours
The H.M.E. must be supplied individually packaged as clinically clean
The H.M.E not impact on performance of the Ventilator breathing circuit
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality ventilator or breathing circuit
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 35 - Anaesthetic Machine end Bacterial Viral Filter for Theatre Use | Must NOT Have a Gas sampling port
Must have smooth edges and sealed housing
Filter Efficiency to be Greater than > 99.9%
The Bacterial/Viral Filter must have a Max time of 24 hours
The Bacterial/Viral Filter H.M.E. must be supplied individually packaged as clinically clean
The connection of Sampling line to Bacterial Viral Filter must be secure
The connection of the sampling line to the filter results in minimal kinking of the line and has no impact on patient end
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the Anaesthetic machine/ventilator or breathing circuit
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 36 - HEPA Bacterial Viral Filter (for use in ICU and or Home Care) | Must NOT Have a Gas sampling port
Must have smooth edges and sealed housing
Filter Efficiency to be greater than 99.99%
The Bacterial/Viral Filter must have a minimum time of 24 hours or greater
The Bacterial/Viral Filter must be supplied individually packaged as clinically clean
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the machine/ventilator or breathing circuit.
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 37 - Adult Manual Resuscitator | Must include Adult Face Mask
Oxygen tubing must be at least 7ft to connect to wall flow meter
O2 Reservoir Bag as part of the Self-inflating resuscitation bag
Air intake valve
Non-rebreathing valves
Standard tubing connecters for attaching to masks or tubes
Must be MRI Compatible. Evidence must be provided.
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 38 - Paediatric Manual Resuscitator | Must include Toddler and Infant Face Mask
Oxygen tubing must be at least 7ft to connect to wall flow meter
Contents must be in a sealed bag
O2 Reservoir Bag as part of the Self-inflating resuscitation bag
Air intake valve
Non-rebreathing valves
Standard tubing connecters for attaching to masks or tubes
Must be MRI Compatible. Evidence must be provided.
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 39 - Expiratory Valve for ICU Ventilator | Expiratory valve
The size/shape of the filter should be deemed not to unduly interfere or impact on the performance or functionality of the ventilator or breathing circuit.
Expiratory valve must fit Draeger ICU ventilator
Standard tubing connecter
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 40 - Oxygen Facemask Medium Concentration with C02 Monitoring | Must come with Tubing
Paediatric and Adult Size
Must not completely cover the face when on the patient must openings around the mask
Reduce the risk of CO2 rebreathing during exhalation,
The mask openings allow expired carbon dioxide to escape
Must have the option to have ETCO2 monitoring available
Must be able to span a range 24-90%
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 41 - Incentive Spirometers | Mouthpiece
Tubing
Resistance chamber
Must be individually packaged
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided
Lot 42 - Spirometers One Way Valve | Mouthpiece
Tubing must be Kink Resistant
One way valve to prevent exhalation
Coaching indicator
Patient goal indicator for supplemental oxygen
Conformity Assessment Markings relevant for N.I
Free from Natural Rubber Latex
Data Sheet Sheets/Technical Information to be provided
Instructions for Use to be provided

## Ethics - Goods
Preliminary Market Engagement Questionnaire for Ethical Procurement
Framework Number | 6012040
Framework Title | Oxygen Therapy Products and Anaesthetic Breathing Systems
Note to Suppliers: | In support of Procurement Policy Note (PPN) 05/21 – Human Rights in Public Procurement and in light of the subject matter of the Contract, the Clients are seeking to gather information to complete a risk-rating exercise with regards to Ethical Procurement. To assist the Clients in the completion of a risk-rating exercise, you are encouraged to complete the questions detailed within: - • Section 1 which focuses on the practice of your organisation as a whole • Section 2 which focuses on supply chain practices specific to (insert goods/services being tendered for)
Questions |  | Viamed Ltd. Response
Section 1: Company Practices
1 | What industry(ies) do you currently operate in and do you see this as being high / medium / low risk with regards to the issue of Human Rights, Modern Slavery and Ethical Procurement? | We operate in the medical devices and healthcare equipment industry. Risk is low–medium: low in Tier 1 supply chains (Germany/USA with ISO-certified suppliers), but medium where occasional sourcing from higher-risk regions may occur. We mitigate through policies, supplier due diligence, and staff training.
2 | On a scale of 1-10 (1 = very weak, 10 = very strong) how would you rate your organisation's general awareness of Ethical Procurement, Human Rights and Modern Slavery? | 7/10 – We have a published Modern Slavery Statement, policies, annual supplier reviews, and staff training in place. Awareness is good for an SME, with improvements ongoing.
3 | On a scale of 1-10 (1 = very weak, 10 = very strong) how would you rate your organisation's ability to carry out due diligence on Ethical Procurement, Human Rights and Modern Slavery? | 7 – We have formal policies, training, and annual reviews in place, with improvements ongoing.
4 | Do you have a Human Rights or Modern Slavery policy or strategy? If no, why not and is this something you intend to create? | Yes – Modern Slavery Statement and Anti-Slavery Policy are published and reviewed annually
5 | Do you adopt any ethical procurement practices? If Yes, do you utilise a remediation-based approach? If No, please explain why not and if you have any future plans in this area | Yes – we assess suppliers’ ethical compliance, request their policies, and are implementing supply chain due diligence
6 | Are you aware of, or do you currently use, the MSAT tool from the UK Home Office? | Yes – we are aware of the ETI Base Code and align with its principles (no child labour, no forced labour, fair working conditions).
7 | Do you have a Supplier Code of Conduct? If no, how do you ensure your suppliers behave responsibly? | In development – supplier terms require compliance with ethical labour standards; formal Supplier Code will be introduced as part of ongoing improvements.
8 | Do you partake in Independent 3rd Party Social Audits? | Not yet – as a small SME, we have not commissioned external audits, but we review supplier ISO13485/ISO9001 certifications annually
9 | Are you part of a global or international supply chain network, and / or do you source goods / services from any countries classified as medium - high risk according to the Global Slavery Index? | No – not currently a member of global audit platforms, but we align our policies with NHS Evergreen and UK Government guidance.
10 | Please answer if question 9 applies and you are part of a global or international supply chain: Where do you sit within the end-end sourcing activity? Where geographically does the supply chain to which you belong originate?
11 | If any, what due-diligence checks do you currently carry-out to prohibit, or prevent unethical practices and / or Modern Slavery within your supply chain? | Annual supplier reviews, request for modern slavery and quality certifications, monitoring via our Intrastats system, internal reporting and escalation mechanisms
12 | Do you and/or your Consortia partners monitor sub-contracting parties within your supply chain(s)? If no, why not? | Yes – through supplier audits (document checks), annual reviews, and by requiring suppliers to provide evidence of compliance.
13 | Are you involved with any NGOs or other similar organisations? | Not directly, but we reference guidance from ETI, CIPS, and UK Government and will consider collaboration in future.
14 | Have you factored in Ethical Procurement, Human Rights and Modern Slavery to your long term business planning? If so what are your longer term aspirations? How do you intend to deliver this? | Yes – incorporated into supplier approval, due diligence, and internal staff training
15 | Do you currently have a supply chain map? | Partial – Tier 1 suppliers are mapped; Tier 2 and below will be extended in future
Section 2: Supply Chain Practices
Supply Chain Geography - Part 1
1 | Are your products commonly manufactured, grown or sourced from or in any of the High Risk countries in Table 1 to the right? | No – main sourcing from Germany, USA, occasional supply from China | Table 1
2 | Are your products commonly manufactured, grown or sourced from or in any of the High Risk countries in Table 2 to the right? | No – no direct raw material sourcing, only finished medical devices. | Highest Risk Countries - All Categories
Supply Chain Geography - Part 2 |  |  | North Korea | All countries in Africa, except South Africa, Botswana, Namibia, Algeria, Tunisia, Morocco, Mali and Senegal | Afghanistan
3 | Are substantial amounts of activity in the supply chain undertaken where little or no effective labour standards legislation exists? | No | Eritrea | Mongolia | Mauritania
4 | Are substantial amounts of activity in the supply chain undertaken where the enforcement of labour standards legislation is weak or non-existent? | No | Cambodia | Turkmenistan | Myanmar/Burma
5 | Are substantial amounts of activity in the supply chain undertaken where there are high levels of poverty and deprivation? | No | Iran | Papua New Guinea | Pakistan
6 | Are substantial amounts of activity in the supply chain undertaken where there is high unemployment? | No
7 | Are substantial amounts of activity in the supply chain undertaken where little or no employee representation (e.g. trades unions) exists? | No | Table 2
Supply Chain Labour |  |  | High Risk Countries by Category
8 | Are substantial levels of labour within the supply chain drawn from more vulnerable groups (e.g. migrants, ethnic minorities, poorly educated, women, convicted)? | No – our direct suppliers are mainly EU/USA based. | Medical Devices/ Pharmaceuticals | Textiles | Food
9 | Are substantial levels of manual labour used within the supply chain? | No – majority are ISO-certified medical manufacturers with automation. | Malaysia | Pakistan | Guatemala
10 | Are substantial levels of low skilled labour used within the supply chain? | No – primarily skilled, regulated workforce in EU/USA | Pakistan | Bangladesh | India
11 | Are substantial levels of seasonal or casual labour used within the supply chain? | No – not applicable to our supply chain. | Afghanistan | Vietnam, Thailand, Cambodia, and Laos | South America - Belize, Columbia, Peru, Brazil, Dominican Republic, Ecuador, Venezuela, Guyana
12 | Are payment cultures based on piece work, cash payment or in-kind payment common place within the supply chain? | No – all suppliers operate under formal, regulated payment structures. | Thailand | Malaysia | Brazil
Supply Chain Transparency |  |  |  | India | East Asia - Seafood
13 | Are there more than two supply chain tiers between you and the product manufacturer/producer? | Yes – for some products, but we engage only with Tier 1 suppliers and are extending due diligence deeper |  | China - Cotton only | Southern and Southeast Asia - Poultry, Seafood
|  |  |  | Indonesia | Spain and Southern Italy - Fruit & Veg/ Non-Meat Food items

## Sustainability & Net Zero
Preliminary Market Engagement Questionnaire - Sustainable Procurement
Framework Number | 6012040
Framework Title | Oxygen Therapy Products and Anaesthetic Breathing Systems
Note to Suppliers: | In support of NI Climate Change Act 2022 & Procurement Policy Note (PPN) 06/21 – Taking account of Carbon Reduction Plans in the procurement of major government contracts, and in light of the subject matter of the Contract, the Clients are seeking to gather information regarding the supply chains' progress to Net Zero. PaLS wish to inform the market that it will soon be a prerequisite for suppliers wishing to bid for all PaLS contracts to confirm your commitment to achieving Net Zero by 2050. Furthermore, for contracts in excess of £5million per annuum and for all frameworks, suppliers may be required to have published a Carbon Reduction Plan detailing clear targets for reducing GHG emissions. Please answer all questions listed in this tab.
Useful Links
NI Climate Change Act 2022 | Climate Change Act (Northern Ireland) 2022 (legislation.gov.uk)
Procurement Policy Note 06/21: Taking account of Carbon Reduction Plans in the procurement of major government contracts | https://www.gov.uk/government/publications/procurement-policy-note-0621-taking-account-of-carbon-reduction-plans-in-the-procurement-of-major-government-contracts
Evergreen Sustainable Supplier Assessment | We would particularly direct your attention to the Evergreen assessment, which will be in use across the NHS: https://www.england.nhs.uk/nhs-commercial/central-commercial-function-ccf/evergreen/
NHS England Supplier Webinars on Net Zero and Carbon Reduction Plans | https://www.events.england.nhs.uk/net-zero-and-sustainable-procurement-team-webinar-programme
Questions |  | Viamed Ltd. Response
Is your organisation committed to achieving Net Zero emissions? |  | Yes – committed to Net Zero by 2050, with interim 2030 targets.
If yes to above, what is your Net Zero target year? |  | 2050 (Net Zero); 100% reduction/offset of Scope 1 & 2 by 2030.
If available please provide a link to your organisations Net Zero Commitment published on your website. |  | https://www.viamed.co.uk (Carbon Reduction Plan published in Policies on website). https://cdn.shopify.com/s/files/1/0380/8149/5085/files/Viamed_Ltd_Carbon_Reduction_Plan_Statement_CRP_Aiming_for_Net_Zero_ID146695.pdf?v=1750321890
Has your organisation completed the NHS England: Evergreen Sustainable Supplier Assessment? |  | Evergreen completed October 2025
If so what was your organisations maturity score? |  | Awaiting maturity score confirmation.
Has your organisation published a Carbon Reduction Plan (CRP) as set out in PPN 06/21. |  | Yes – Carbon Reduction Plan published and aligned with PPN 06/21.
If available, please provide a link to your Carbon Reduction Plan published on your website. |  | https://www.viamed.co.uk (Carbon Reduction Plan available in the Policies section on the website). https://www.viamed.online/pages/policies