Why the VM-2500 does not require calibration 

Version 1.1 

August 2018

	 [image: image1.png]



	VIAMED
	
	Viamed Limited - 15 Station Road - Cross Hills

Keighley -West Yorkshire  BD20 7DT - United Kingdom
Tel: +44 (0)1535 634542     Fax: +44 (0)1535 635582

Email: info@viamed.co.uk   Website: www.viamed.co.uk


Why the VM-2500 does not require calibration during use
The CO2 analyzers of VM-2500 Series capnographs (mainstream and sidestream) are permanently factory calibrated during the manufacturing process. During extensive measurements with calibration gas, the CO2 measurement units evidenced the accuracy required. The monitor requires no routine calibration and the CO2 analyzers incorporate maintenance and calibration-free technology

All necessary calibration constants are stored within each CO2 analyzer, the analyzer firmware contains algorithms that self-compensate for any drift over time caused by changes from the original component specification (e.g. changes within the optics, alteration of the light transmission properties of the windows in the mainstream airway adapter etc.) and self-adjusts its settings based on the information that it generates during start-up and self-test procedures. The use of Viamed consumables used in accordance with the user manual enables the accuracy and function of the analyzers to be maintained.
The analyzers have been rigorously tested and proven to consistently give accurate results without the need to verify accuracy against a reference gas, which is a major advantage of this unit over many alternative capnographs that need: recalibrating, accuracy verifying or servicing on a regular basis.
Should it be required by local ordinances, there is an option within the service menu to verify the measurement accuracy using a suitable tight-tolerance, medical grade reference gas mixture. Alternatively, CO2 gas readings can also be assessed at regular intervals against a reference instrument.

A basic maintenance plan conducted by service personnel is recommended. Should you not have personnel to carry this out, capnographs can be returned to Viamed for servicing. Please refer to the user manual and service manual for detailed information. 
Should you have any queries or require further information, please do not hesitate to contact Viamed or your local distributor.
