Hi Jerry,
The above is our reference to the problem.

You will probably have received my latest email to him.
I suggest you take the two older sensors back and we will replace them.
Our overall failure rate is still (<0.15%)  which is very low.
However we are seeing some returns with high output.
As I explained in DEMA sensors can fail low due to lead depletion of high due to evaporation. These sensors appear to be failing high.
The test rig tests are confusing so we really need them back here as soon as possible.

Now that we have isolated this infrequent effect we are working on a solution and will inform you as soon as possible.

Just out of interest and to add to my education.
What tests do you recommend the diver does underwater if he/she feels they may have a sensor problem. I have always recommended a test with Oxygen at 6mtrs and a diluent test at depth.
These should show possible problems with current limiting or deviation from the set point.

John