oxygen monitoring

Oxygen Analyser and Monitor

AX300 oxygen analyser and MX300 oxygen monitor

Ideally suited for use in Respiratory Therapy, Neonatal and Anaesthesia applications

The microprocessor based AX300 Oxygen Analyser and MX300 Oxygen Monitor offer a range of unique standard features that set these models apart from the competition.

Featuring easy calibration and settings which are secured automatically. In addition the low power requirement allows the units to operate continuously for 2000 hours on just 3 x AA alkaline batteries. Both models also feature the R-17MED oxygen sensor providing superior long-term stability and accuracy.

Constructed from high impact resistant materials, these rugged instruments provide unparalleled performance under the harshest conditions.

Teledyne Oxygen Monitors 
and Analysers

AX300 Analyser features
Microprocessor controlled for 


enhanced performance

Large easily read backlit LCD display

Auto calibration

18% Hypoxia alarm

Battery strength indicator

Auto diagnostics with fault codes

Sensor disconnect/failure alarm

R-17MED oxygen sensor with flow diverter

Sensor cable

Mounting stand

‘V’ mount adapter

22mm ‘T’ adapter

MX300 Monitor additional features
Two fully adjustable alarm set points


(range 18 - 100%)

Alarm test key

Alarm silence key

Alarm OFF function

Battery test key

	
	

	Analysis range
	0 - 100 % Oxygen

	Full scale accuracy
	± 2% at constant temperature and pressure

	Response time
	90% of step change < 6 seconds at 25°C

	Operating temperature
	0 - 40°C

	Storage temperature
	10 - 30°C (continuous), 5 - 50°C (intermittent)

	Power requirement
	3 x AA alkaline batteries

	Battery life
	Minimum 2000 hours in no-alarm, 
no output signal state

	Sensor type
	Teledyne class R-17MED

	Sensor life
	36 months in air 
(approximately 7.5 months in 100% O2)

	Sensor cable
	Coiled, extendable

	Weight
	420g

	Dimensions with stand
	140 mm (H) x 88 mm (W) x 76 mm (D)

	Signal output
	0 - 1 VDC or RS-232 (2400 Baud rate)

	RFI hardened
	Yes

	Compliance
	ASTM F 1462 ASTM F 1463 ISO 7767 ISO 9703-1 ISO 9703-2 EN/IEC 60601-1 EN/IEC 60601-102 MIL-STD-810E

	Regulatory approvals
	CE, CSA and FDA


Terms Explained...

Outputs

Analogue

RS232

Actual Size Screens

2.5 Digits

3.5 Digits

Includes batteries, operator manual and the following standard accessories:

Sensor Cable 
(Coiled 0.9m, extendable) Right angled connectors

R-17MED 
Oxygen Sensor

Sensor Flow Divertor

15mm O.D.

‘T’ Adapter

22mm I.D. - 22mm O.D., 
T = 15mm I.D.

Mounting Stand

‘V’ Mount Adapter

See back page for optional extra accessories...

Oxygen Analysers
with fixed 18% low alarm


Standard Unit
3.5 Digit Display

Analogue Output
3.5 Digit Display

RS232 Output
2.5 Digit Display

Analogue Output
2.5 Digit Display

RS232 Output
Actual Size
3.5 digit display

Oxygen Monitors

with adjustable low & high alarms
3.5 Digit Display

Analogue Output
3.5 Digit Display

RS232 Output
2.5 Digit Display

Analogue Output
2.5 Digit Display

RS232 Output
Actual Size
3.5 digit display

Teledyne Oxygen Monitors 
and Analysers

Additional Optional Accessories

Sensor Cable 
(Straight 3m) 
Right angled connectors

Sensor Cable 
(Coiled 1m, extendable) Straight connectors

Interface Cable (1.5m) for analogue output 

Interface Cable (1.5m) for RS232 output 

Sensor Adapter Cap 22mm I.D.

Sensor Adapter Cap 
22mm O.D.

Oxygen Concentrator Adapter Cap

Callibration Assembly

Perfusion Circuit 
‘T’ Adapter Set

‘T’ Adapter 15mm I.D. - 15mm O.D., T = 15mm I.D. for paediatric circuits

‘T’ Adapter 
22mm I.D. - 22mm O.D., T = 15mm I.D./22mm O.D. Extended ‘T’

‘T’ Adapter 
22mm I.D. - 22mm O.D., 
T = 15mm I.D./22mm O.D. 
Non re-breathing

‘T’ Adapter 15mm O.D. / 22mm O.D. - 22mm O.D., 
T = 15mm I.D./22mm O.D. 

Non re-breathing

‘T’ Adapter 
22mm I.D. - 22mm O.D., 
T = 15mm I.D./22mm O.D.

‘T’ Adapter 
22mm O.D. - 22mm O.D., 
T = 15mm I.D./22mm O.D.

‘T’ Adapter  - Metal
22mm I.D. - 22mm O.D., 
T = 15mm I.D.

Connector 
15mm O.D. - tapered barb

Connector 15 mm I.D./ 
22mm O.D. - tapered barb

Connector 22mm I.D./ 
30mm O.D. - tapered barb

FiO2 Sampling System
22mm I.D. - 22mm O.D.

Replacement ‘T’ Adapter 
for 0120164

Replacement Activator for 0120164

‘V’ Wall Mount 

‘V’ Mount Pole Clamp

Small: 
14 - 25mm clamping width

‘V’ Mount Pole Clamp

Large:
16 - 40mm clamping width

‘V’ Mount Pole Clamp Horizontal

10 - 40mm clamping width

‘V’ Mount Rail Clamp
Attaches to most types of medirail - 30mm (H) x 10mm (D)

Table Top Swivel Mount*

Universal Mounting Clamp*

* Attaches to the threaded socket on the rear of the units

Specifications Subject to Change

Distributed by:

Part Number:

Date:

