| Measurement Range: | 0-100 % oxygen |
| --- | --- |
| Output in ambient air: | 9 to 13mV |
| Electrical Interface: | OOD103: 3pin (Molex 22-11-1031) |
|  | OOD103-1: 3.5mm phone jack (Switchcraft type 41) |
|  | OOD103-1V: Mini HF (RADIALL R114.426.000) |
|  | OOD103-2V: Mini HF (RADIALL R114.186.000) |
|  | OOD103-OG: 3pin (Molex 22-11-1031) |
|  | OOD103-ISC: 3pin (Molex 22-11-1031) |
| Repeatability: | < 1 % vol. O2 @ constant temp. and pressure |
| Linearity error: | 0.1 to 1.0 bar less than 3 % relative |
|  | 0.1 to 2.5 bar less than 5 % relative |
| Response time: | < 5sec. to 90 % of final value |
| Zero Offset Voltage: | < 0.5% oxygen reading in 100 % nitrogen @ 25°C after 36 seconds |
| Cross Interference: | < 0.5 % vol. O2 response to: |
|  | 5 % CO2 balance N2 |
| Influence of Humidity: | - 0.03% rel. per % RH at 25°C |
| Pressure Range | 600 to 2500 mbar |
| Influence of Mechanical Shock: | < 1% relative after a fall from 1m |
| Operating Temperature: | 0 to 50°C |
| Temperature Compensation: | built-in NTC compensation |
| Effect of Temperature Compensation: | between 0 °C and +50 °C: 5 % relative error (steady state) |
| Operating Humidity: | 0-99% RH non-condensing |
| Long Term Output Drift: | < 1% vol. oxygen per month |
|  | typically < - 15 % relative over lifetime |
| Storage Temperature: | -20 to +50 °C |
| Recommended Storage: | +5 to +15 °C |
| Recommended Load: | 10 kOhms |
| Warm-Up Time: | < 30 minutes, after replacement of sensor |
| Nominal Sensor Lifetime: | 500,000 % vol oxygen hours |
| Weight: | approximately 28 grams |
| Warranty Period: | 15 months (including 3 months shelf life) |
| Part No.: | OOD103: 1001570 (REF 8010000 / VST2200) |
|  | OOD103: 1001570 (REF 8010001 / APD 5) |
|  | OOD103: 1001570 (REF 8010006 / CR22D) |
|  | OOD103: 1001570 (REF 8010008 / NaNM01) |
|  | OOD103: 1001570 (REF 8010014 / K-22D) |
|  | OOD103: 1001570 (REF 8010016 / H22) |
|  | OOD103: 1001570 (REF 8010015 / UDT-22) |
|  | OOD103: 1001570 (REF 8010018 / M3S) |
|  | OOD103-1: 1001580 (REF 8010003 / VR1700) |
|  | OOD103-1: 1001580 (REF 8010005 / VST1700) |
|  | OOD103-1: 1001580 (REF 8010017 / VMS1700) |
|  | OOD103-1V: 1001581 (REF 8010002 / APD 11) |
|  | OOD103-1V: 1001581 (REF 8010007 / NaNS01) |
|  | OOD103-1V: 1001581 (REF 8010022 / APD 16) |
|  | OOD103-2V: 1001721 (REF 8010004 / JJ-CCR) |
|  | OOD103-OG: 1001930 (REF 8010010 / RBL-23D) |
|  | OOD103-ISC: 1002106 (REF 8010012 / SC-I22D) |
| All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air |  |

All dimensions in mm

Customized Design Specification:

- Integration of additional Zitex membrane in sensor gas diffusion way

- 2mm hole drilled in sensor cap near to the connector at sensor type OOD103-D

- Conformal coating on the sensor PCB with PLASTIK 70

- 2 x 2mm holes drilled in the sensor PCB

- None information about EnviteC on PCB top side / bottom side only information for ordering

process and traceability

PCB Layout:

Type:  8010000 / 8010001 / 8010003 / 8010004 / 8010005		    8010002 / 8010007 / 8010022

/ 8010006 / 8010008 / 8010010 / 8010012 / 8010014

/ 8010015 / 8010016 / 8010017 / 8010018

Top			    Bottom			         		  Top			           Bottom

Marking of product

Customized Label Specification is part of separate document.

ENVITEC- WISMAR GMBH

Alter Holzhafen 18

D-23966 Wismar

Phone:  +49-(0) 3841- 360 1

Fax:      +49-(0) 3841- 360 222

E-Mail: info@envitec.com

http://www.envitec.com