| Measurement Range: | 0 % ... 100 % oxygen (at atmospheric pressure) |
| --- | --- |
| Output in Ambient Air: | 9 mV ... 13 mV |
| Electrical Interface: | 3 pin (Molex® 22-11-1031) |
| Repeatability: | < 1 % volume O2 at constant temperature and pressure |
| Linearity error: | < 3 % relative |
| Response time: | < 12 s to 90 % of final value |
| Zero Offset Voltage: | < 200 μV in 100 % nitrogen, applied for 5 min |
| Cross Interference: | meets ISO 80601-2-55 requirements |
| Influence of Humidity: | -0.03 % rel. per % RH at 25 °C |
| Pressure Range: | 0.6 bar ... 2 bar (ppO2 0 ... 1250 mbar O2) |
| Influence of Mechanical Shock: | < 1 % relative after a fall from 1 m |
| Operating Temperature: | 0 °C ... +50 °C |
| Temperature Compensation: | built-in NTC compensation |
| Effect of Temperature Compensation (steady state): | between +25 °C and +40 °C: 3 % relative error between 0 °C and +50 °C: 8 % relative error |
| Operating Humidity: | 0 % ... 99 % RH non-condensing |
| Long Term Output Drift: | < 1 % volume oxygen per month |
|  | typically < -15 % relative over lifetime |
| Storage Temperature: | -20 °C ... +50 °C |
| Recommended Storage: | +5 °C … +15 °C |
| Recommended Load: | ≥ 10 kOhms |
| Warm-Up Time: | < 30 minutes, after replacement of sensor |
| Nominal Sensor Lifetime: | ≥ 1 000 000 % volume oxygen hours |
| Weight: | approximately 28 grams |
| Warranty Period: | 15 months (including 3-month shelf life) |
| Part No.: | OOM102-A: 1002504 (REF 0110039 / R-22VA) |
| All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air. |  |

Mechanical draft:

General tolerances ISO 2768-c

*Intermediate thread: Metric / Unified Extra Fine

Customized Design Specification:

- NA

Product Labeling

| OOM102-A: 1002504 (REF 0110039 / R-22VA) |  |
| --- | --- |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |

Customer Approval

Date:

Signature: