	Subject
	:
	faulty sensors
	

	Date
	:
	Tue, 12 May 2015 10:20:00 +0100
	

	Linked to
	:
	Jessica Wagner
	

	From
	:
	"Wagner, Jessica (GE0Y)" <Jessica.Wagner@Honeywell.com> (By way of info@viamed.co.uk)
	

	To
	:
	SNIXON (Steve Nixon) <GoldMine User>
	



Hi Steve,

Here the conclusion of our meeting written by Christa:

The problem is that the membrane is not in close contact with the electrode as required. The membrane was displaced after mounting by strong vibrations in a subsequent production process. A relaxation process over month takes place leading to a slow drift in signal over the first months after production (up to 1mV /month of signal in air). Furthermore in some cases the offset rises and the sensor signal changes if the sensor is moved around. 

We realized the long term problems only after your complaint last week. We checked our retained samples and found that a high percentage of the sensors are drifting. Currently most of the retained samples are inside our specification. It is difficult to predict how many sensors will drift outside the specification. 

The problem is probably accelerated by non stationary applications and pressure changes. 

The following serial numbers of OOD sensors are affected: 

	Serial number.
	number

	
 
	
 

	
VMS 100421-100500
	80

	
JJ 105924-106123
	200

	
ISC 105901-106000
	100

	
S01 105501-105600
	100

	
ISC 106001-106100
	100

	
S01 105601-105700
	100

	
M01 102651-102750
	100

	
JJ 106124-106323
	200

	
APD 106751-107000
	250

	
503 17232-17381
	150

	
K 100556-100605
	50

	
S01 105701-105800
	100

	
M01 102751-102850
	100

	
503 17382-17531
	150

	
JJ 106324-106523
	200


For the yellow marked serial number it is not clear whether they are affected. 

We propose a telephone conference to discuss further procedure. 

We have a question. How often are the sensors calibrated? 

Kind regards / Mit freundlichen Grüßen

i.A. Jessica Wagner

Country Manager

EnviteC-Wismar GmbH a Honeywell Company
Alter Holzhafen 18
23966 Wismar, Germany 
www.envitec.com



Phone: +49 3841 360 206

Fax: +49 3841 360 222

Mobile: +49 172 5211756

Jessica.Wagner@honeywell.com
Registergericht: Schwerin, HRB 1838
Geschäftsführung: Kolja Kress, Marcus Ostländer

