VIAMED EVALUATION FREETEXT COMMENTS

Good

Can leave without worrying about weight on tracheostomy

Easy to have on the bed

Easy to place-use

No weight on tracheostomy, well away from secretions

Easy to use

Very useful to have by anaesthetist as intubating

Anaesthetist liked it: had all obs needed and audible desaturation alert

In CT waveform visible behind the glass

In ICU able to transfer to bed off big monitor and still have CO2 RR and sats

Batteries low when turned on, lasted approx. 1 hour after being changed

Alarm not picked up on so device cut out

Radiographer commented on the lack of slack for CT scan (there was enough)

Unable to see through glass in CT

Tried to use on NIPPV, inaccurate readings

Immediate reading

Excellent

ICU Reg would like a finapres for BP

Transferred with patient on IABP: only monitoring equipment needed

Batteries running out when only recently charged: difficult to see as battery sign small

Good as CPR feedback device

Lithium battery worked

Size of machine great at head of bed

Very easy to use