[image: image1.jpg]West Midlands Ambulance Service m

NHS Foundation Trust






West Midlands Ambulance Service NHS Foundation Trust

Specification for

End Tidal CO2 Capnography Device
Tender Reference 0131-14-W
Contents Page

Contents 






   Page No.

Contents Page






2

1.0
 Introduction





3
2.0 
Requirement





3
3.0
 Specification





3
4.0 
Sample






5
5.0
 Delivery






5
6.0 
Contract Award





5
7.0 
Terms & Conditions




6
1.0 Introduction
West Midlands Ambulance Service NHS Foundation Trust is procuring End Tidal CO2 Capnography Devices. Upon completion of the procurement process the contract will allow all UK ambulance services to procure from it as they require. However, there will be no level of guaranteed business from this. 

Upon award of contract the contract will be open to all UK Ambulances to procure from it.
2.0 Requirement

As a direct result of a national regulation 28 Coroners order West Midlands Ambulance Service NHS Foundation Trust is introducing End Tidal Carbon Dioxide monitoring to all vehicles with immediate effect.  The devices need to have both Capnography and Capnometery on them and have minimal tubes linked to them to assist when carrying out manual handling of patients. 

The Trust will utilize the device for all ventilated adult and pediatric patients irrespective of airway device utilized e.g. ET or Supraglotic.
This initial order requirement will be for 450 devices plus 900 adult and 900 pediatric to enable two sets of each to be stocked with each device. In addition to this there will be a requirement of circa 4000 tubes per annum inline with current activity levels (this value may increase or decrease subject to activity levels during the life time of this contract).
3.0 Specification 

The Trust requires the device to be fully portable with Capnography and Capanometery on them and have minimal tubes linked to them to assist when carrying out manual handling of patient and capable of meeting all of the detailed specification requirements to enable the device to be practically used within the field of pre-hospital care.
	
	Requirement

	Measurement device
	Non-dispersive IR absorption

	Monitoring Range
	CO2  


0-99mmhg or greater




0-9.9kPa or greater

Respiratory rate =
3-150bpm

	CO2 Accuracy Range
	No more than 8% of the reading at 99%

	Respiratory Rate Accuracy Range
	+/- 1 respiration per minute

	Operating Temperature
	-5 - +50°C (23 to 122°F)

	Operating Atmospheric pressure
	70 - 120 kPa

	Operating Humidity
	< 40 hPa H2O

	Temperature Range
	-30 - +70°C

	Storage atmospheric pressure
	50 - 120 kPa

	CE
	93/42/EEC Medical Device Directive

	IP
	>IP32

	Warm up time
	<20 seconds

	Calibration
	The unit must have no routine calibration requirements due to its operating environment

	Single patient use adaptor
	Mandatory requirement

	Cables and tubes
	The device should have a maximum of one tube and one cable

	Weight
	<100grams

	Life expectance
	>2 years

	Consumable requirements
	Adaptor

Tube and adaptor

	Power source
	Lithium batteries

Rechargeable

Disposable

	Product casing 
	Carry case

Consumable portable case

	Cost Options
	Single unit -

<50  or  50-100

100-200 or 200-300   

<300


3.1 Training Requirements:

Theoretical training pack to be provided, to include training material such as user guide with the choice of in house or external training to be provided in conjunction with the awarded supplier. 

Along with a training card to be issued with each device or member of staff.
3.2 Support, Maintenance & Warranty:
Manufacturer and/or supplier standard warranties would apply, further details are asked to be provided within the Technical submission.
4.0 Samples 

One sample of the product you would intend to supply in line with the specification will need to be provided, the sample must reflect and match the product being sold.

The sample needs to be delivered to the following addressee by Friday 10th October 2014

Mrs Amarpreet Kooner, 

WMAS Procurement, 

Chantry House, 

High Street, 

Coleshill, 

Birmingham, 

B46 3BP 
5.0 Delivery:

Due to the Coroners regulation 28 publication to all Ambulance Services time is of the essence for this provision once the contract has been awarded.
6.0 Contract Award
The contract will be awarded for a 3 year term with the option to extend the contract for a further 2 years at two 12 monthly increments

Any quantities stated within the invitation to Tender are indicative only West Midlands Ambulance Service NHS Trust cannot guarantee any level of business arising from the award of this contract.
The contract is open to all Ambulance Services across England, Scotland, Northern Ireland and Wales to purchase from. UK ambulances services are (correct as of September 2014):

· North East Ambulance Service NHS Foundation Trust

· North West Ambulance Service NHS Trust

· Yorkshire Ambulance Service NHS Trust

· West Midlands Ambulance Service NHS Foundation Trust

· East Midlands Ambulance Service NHS Trust

· East of England Ambulance Service NHS Trust

· South Western Ambulance Service NHS Foundation Trust

· South Central Ambulance Service NHS Foundation Trust

· South East Coast Ambulance Service NHS Foundation trust

· London Ambulance Service NHS Foundation Trust

· Welsh Ambulance Service

· Scottish Ambulance Service

· Northern Ireland Ambulance Service

7.0 Terms and Conditions

The NHS Terms and Conditions for the Supply of Goods will govern this contract for entire lifetime of the contract. A copy of the NHS Terms and Condition for the Supply of Goods can be found on the attachments area of the Tender portal. 

Specification – 0131-14-W End Tidal CO2 Capnography Device


6

[image: image1.jpg]