QC 21 Non Conformance Report

Date	25/08/15
Issue id	61743
BSI Ref (if applicable)	1225836M1
Responsibility	Derek Lamb
Non-Conformance	1225836M1 / “The elements for the clinical evaluation are all available in different places in the technical file but the requirement for evaluation of this data has not been carried out as required by Annex X of the MDD. There is evidence from 2010 that the device was recommended for use by the NLS and the UK Resuscitation council. “
Investigation By:	Derek Lamb
Investigation Issue id / (if applicable)	The post market surveillance and risk assessment procedure requires updating so a formal follow-up document is produced to show the “evaluation of the data”
Corrective Action By:	Derek Lamb
Corrective Action Issue ID (if applicable):
Time Scale for Corrective Action
Corrective Action:	Added/updated procedures in the Viamed ISO system / VM3COP27.11 Performing a Technical File PMS and risk assessment and VM3COP27.12 Clinical Evaluation Risk assessment Technical Files The Process of doing the yearly PMS / Risk assessment now directly routes to a final clinical re-evaluation and risk benefit against residual risks document being produced and added to the files. / The risk of fire has been added to the core risk analysis questions – which has triggered a requirement to re-review all the technical files.
Follow-up future issue id / (Effectiveness verification)
Effectiveness verification
Closed By:
Closed on