QC 21 Non Conformance Report

Date	25/08/15
Issue id	61746
BSI Ref (if applicable)	1225836N3
Responsibility	Derek Lamb
Non-Conformance	“The risk management plan is a live system which forces responsibility to be taken by the manufacturer / through the issue system. There are a couple of elements missing from the 14971:2012 standard. There is / no requirement that residual risks are disclosed in the instructions for use. A risk benefit analysis is not / performed on residual risks. A multidisciplinary team is not yet required to keep the living document up to / date. The risk of fire has not been covered in the questions associated with the risk analysis. It is, however, / addressed in the choice of materials.”
Investigation By:	Derek Lamb
Investigation Issue id / (if applicable)	Final Post market surveillance and risk assessment reports require updating to add a final summary report, incorporating the risk/benefit analysis / Core risk assessment questions has missed the section on fire when the risk assessment system was computerised from the paper files.
Corrective Action By:	Derek Lamb
Corrective Action Issue ID (if applicable):
Time Scale for Corrective Action	All files now require a full risk review VM3COP27.11. Once the Tom Thumb file is reviewed and the Final clinical evaluation and the residual risks against the benefits review has been approved by the BSI scheme manager the rest of the files will be re-reviewed.
Corrective Action:	Added/updated procedures in the Viamed ISO system / VM3COP27.11 Performing a Technical File PMS and risk assessment and VM3COP27.12 Clinical Evaluation Risk assessment Technical Files The Process of doing the yearly PMS / Risk assessment now directly routes to a final clinical re-evaluation and risk benefit against residual risks document being produced and added to the files. / The risk of fire has been added to the core risk analysis questions – which has triggered a requirement to re-review all the technical files.
Follow-up future issue id / (Effectiveness verification)
Effectiveness verification
Closed By:
Closed on