Clinical Evaluation research:

Searched for Clinical papers on

Comparable devices
Fisher & Paykel Neopuff 
Fanem Babypuff 1020 
GE Healthcare Panda T-Piece System 
GE Healthcare Giraffe T-Piece System 
GE Healthcare Giraffe Stand Alone T-Piece System 
Draeger Resuscitaire T-Piece Resuscitation  

T Piece resusitators
Results

· Obstetric and Intrapartum Emergencies: A practical guide to managerment.
 Discussing Airway and Newborn Infants.

 Page 243 Recommending use of a bag and mask or preferably a Tom-Thumb type device.

 Published : 2013, 2013

· Mid Essex Guideline for Neonatal Resusitation
 Appendix D – Required equipment Tom Thumb

 2013

Neonatal resuscitation – a practical

approach. The experience of one UK

tertiary neonatal unit

Document id#15733

· Optimal neonatal ventilation 
http://gradworks.umi.com/34/53/3453856.html
2011
by Shelton, Julie A., D.N.P., THE COLLEGE OF ST. SCHOLASTICA, 2011, 72 pages; 3453856 

The Neopuff Infant Resuscitator recently selected as a standard of care in the SMMC Neonatal Intensive Care Unit (NICU) and Birthplace, is a method of ventilating neonates that consistently delivers the same pre-set pressure with a consistent blend of oxygen and air; thus adequate ventilation/administration is not dependent on caregiver technique and experience. 

 University of Kentucky / UK HealthCare Policy and Procedure RC04-04P

 Recommends usage of Neo Puff (T Piece resusitors) 

 Document ID#15734
Advantages of the Neopuff Resuscitator Operator sets PIP and PEEP pressures Resuscitator will not deliver PIP or PEEP above set pressures PIP and PEEP pressures are displayed on unit Operator controls the length of inspiratory time Resuscitator delivers 100% free flow oxygen reliably  

US National Library of Medicine

http://www.ncbi.nlm.nih.gov/pmc/articles/PMC2675376/
and

A comparison of three neonatal resuscitation devices. 2005

http://www.ncbi.nlm.nih.gov/pubmed/16081202
CONCLUSIONS:
The T-piece resuscitator delivered the desired pressures more accurately, but required greater time to increase the PIP from 20 to 40 cmH2O. It was difficult to maintain a prolonged inflation and deliver the desired PEEP with the self-inflating bag even with the PEEP valve in place. There is a need for improvement in the design and function of current manual resuscitation devices and for prospective trials to evaluate the optimal method of bag and mask ventilation during resuscitation of the newborn infant.
NRP (Neonatal Resuscitation Program) 6th Edition

Lessons in positive pressure ventilation 

includes usage of T Piece resusitators

Manual of Neonatal Care
edited by John P. Cloherty, Eric C. Eichenwald, Anne R. Hanse  
It offers greater control over manual ventilation by delivering breaths of reproducible size (peak and end-expiratory pressure) and a simplified method to control delivered breath rate. 

Using a regulated means of admisniration, such as a T-piece resusciator or ventilator, is preferable.  

WORKSHEET for Evidence-Based Review of Science for Emergency Cardiac Care 

2009
Includes Risks / Benefits of T Piece resuscitator units 
Document id#15731

and

Document id# 15732
Clinical question :


In newborns (P) requiring positive pressure during resuscitation, is positive pressure ventilation by T-piece resuscitator (I) superior to bag ventilation (C) for improving outcome - specify (O)? 

 27 JaN 2010,

CONSENSUS ON SCIENCE: 

• There are no clinical studies in newborns requiring positive pressure during resuscitation to support or refute superiority of the T-piece resuscitator over bag ventilation in improving outcome. 

• In mechanical models, target inflation pressures are delivered more consistently when using T-piece resuscitators than with self-inflating bags. (LOE 5 - Hussey 2004, F490; Oddie 2005, 109; Finer 2001, 299) 

• In mechanical models, PEEP is maintained more consistently with T-piece resuscitators compared with self-inflating bags. (LOE 5 - Finer 2001, 299) 

• In mechanical models, the ability to deliver a sustained inflation is better with either a T-piece resuscitator or flow-inflating bag compared with a self-inflating bag. (LOE 5 - Bennett 2005, 113) 

