QC 21 Non Conformance Report

Date	25/08/15
Issue id	61890 and 61891
BSI Ref (if applicable)	1225837M1 / 1225837N1
Responsibility	Derek Lamb
Non-Conformance	“There is no statement of validation of the packaging to go with the packaging trials. ”
Investigation By:	Derek Lamb
Investigation Issue id / (if applicable)	“Aspects of the technical file which were not available for audit included: Compliance with the Essential Requirements Biocompatibility (V&V) Risk Analysis and Control Summary Product Verification and Validation Clinical evaluation OEM labels for the devices “ / in light of BSI Ref: 1225837N1
Corrective Action By:	Steve Nixon / Envitec to send required documentation to BSI directly. / “There is no indication of the date the device should be used by and it seems appropriate that there should be as these items may be stored prior to use and do not last forever ” / We were about to have redesigned labels for the sensors to more match our exisiting similar products from teledyne, the latest design sign off 31/07/2015 has now been scrapped. No products produced with design docid:???? / We are currently in discussions with the original OEM with regard the used by date.
Corrective Action Issue ID (if applicable):
Time Scale for Corrective Action	Depends on OEM, will chase up this week 26/08/2016
Corrective Action:	Will chase Envitec this week to send the required files. / Unknown at this point if they have been received by BSI
Follow-up future issue id / (Effectiveness verification)
Effectiveness verification
Closed By:
Closed on