	[image: image1.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid





	
	[image: image2.png]


 Ltd.
15 Station Road, Cross Hills, Keighley, West Yorkshire, BD20 7DT, UK.

Website: www.viamed.co.uk.  Email: info@viamed.co.uk.
Tel: +44 (0)1535 634542.  Fax: +44 (0)1535 635582.


.
[image: image1.jpg] Tom Thumb Resuscitation Unit Servicing Manual (TT490).
Version 1.1.  16/07/04



Part no: 0390023.

CERA THERM 600-3 radiant heater
Nufer Medical AG

	Check ID
	Question / Area to Check
	OK
	Comments

	Mechanical condition
	
	

	1
	Check heating element for damage (cracks, chipping, loosening)
	
	

	2
	Check housing and chassis (cracks, deformations)
	
	

	3
	Check front film, warning signs and specification plate
	
	

	4
	Tripod dolly (if present): Check rollers and brakes
	
	

	5
	Tripod dolly (if present): Can be fixed in place, height can be altered
	
	

	6
	Tripod dolly (if present): Check shock absorbers
	
	

	7
	General checks for tripod dolly, wall and ceiling mounts (if

present): Paintwork, cracks in plastic parts, deformations, loose

parts, loosening
	
	

	8
	Wall and ceiling mounts (if present): Check fastening element on unit mount every 24 months: thickness at least 1 .5mm, length at least 11 .5mm (see document 18.002.x, request if needed)
	
	

	Function test
	
	

	9
	Check mains switch (correct functioning and display)
	
	

	10
	Check buttons and displays (correct functioning and display)
	
	

	11
	Measure power consumption at 240 V AC I 50 Hz:

(Power consumption is variable - check the lowest value)

Output level 1

Output level 4
	
	

	12
	Check lighting (both levels)
	
	

	Alarms
	
	

	13
	Acoustic and visual reminder alarm after
	
	

	14
	Alarm reset, reminder alarm
	
	

	15
	Acoustic and visual mains failure alarm after at least
	
	

	16
	Alarm reset, acoustic mains failure alarm
	
	

	Electrical safety
	
	

	17
	Check mains cable and unit socket for damage
	
	

	18
	Electric safety test to standard EN 62353:2008
	
	

	Misc./cleaning
	
	

	19
	Cleaning, inside and outside (remove any dust in the unit)
	
	

	20
	Attach service label with new due date
	
	

	
	
	
	

	The unit has passed the safety checks and can be used for the intended purpose.
	
	


Date: 


Name:



Signature:
[image: image2.png][image: image3.jpg]ARAB HEALTH
2002

Summary of products

Pulse Oximetry Sensors Viamed  Dolphin  Aristo
Finger sensors v v v
Multi-site 'Y’ sensors v v v
Ear sensors v
Extension cables v v v
Adapter cables v v v
D.O.T. Sensors v
Wrap sensors v
Disposables v v ~

Pulse oximetry sensor repair services
Pulse oximeters - Stand alone, hand held, compact PC
Pulse oximetry simulators and testers

Oxygen sensors
Oxygen monitors and analysers

Oxygen hoods
Patient monitors
Infant resuscitation cabinets

Infant resuscitators
Radiant warmers

Phototherapy masks
Phototherapy light shields

ECG cables

Temperature probes

Supra-maximal nerve stimulator

Chart & recorder paper [
VIARMED

Imail infeid




