AlcoTrue® P   product and service training course
AlcoTrue® P
product and service training course
Februar 2011

---
AlcoTrue P – short product information
AlcoTrue P
AlcoTrue® P is intended for quick and accurate determination of alcohol concentration in a	person’s breath
AlcoTrue® P is for mobile use with alkaline and rechargeable Li-ion batteries.
Calibration
AlcoTrue® P is provided with standard calibration by bluepoint MEDICAL GmbH & Co. KG
The period of calibration validity is 6 months (typically), after calibration period expires the	device will require re-calibration
The device can either be re-calibrated by bluepoint MEDICAL GmbH & Co. KG or an	authorized service partner
The Expiry date for calibration is displayed on the “Self test” screen when the device is	switched on

---
AlcoTrue P standard package
AlcoTrue® P device
Mouthpieces (25 pieces individually packed)
Alkaline batteries pack (4x AA, LR6, MN1500)
User manual
Accessories
Printer
Rechargeable Li-ion battery
USB Li-ion battery charger
USB cable

---
AlcoTrue P - Gerät
Display
“ON/OFF/Confirm” button
Back button “Back/Menu”
Print button (Optional function)
Arrow buttons “up/down”
USB Interface
Device inlet
Device outlet
Device label
Battery compartment lid
1
1
7
2
8
3
6
9
4
2
5
3
6
5
10
4
7
8
9
10

---
Insert / exchange batteries
4 x alkaline batteries

---
Insert/exchange Li-ion battery
1 x Li-ion battery

---
Battery level indications
Battery indicator top left
100% battery capacity ; battery full
75% battery capacity
50% battery capacity
25% battery capacity
0% battery capacity ; battery empty
Alarm if battery is empty
A periodic alarm tone and blinking display message <Battery empty>
indicate that the batteries must be exchanged or charged
The device automatically switches off after 10 seconds

---
Charging Li-ion battery
Use USB Li-ion battery charger and USB cable
Charging process is running
“Battery full” indicates that Li-ion
 battery is fully charged

---
Device switch ON / OFF
Keep button pressed for 2 s,
device turns on accompanied
with an acoustic tone (beep)
A self test of the internal
functions and components
is automatically performed
Keep button pressed,
countdown is displayed:
device switches off after 3 s
Software and language
versions of the device are
shown
Calibration expiry date are
shown
Automatic switch off - AlcoTrue® P switches off automatically 3 minutes after the last use.

---
Insert mouthpeice
Hygienic mouthpieces type S ( bluepoint medical)

---
Alcohol level indication
AlcoTrue P- breathalyzer is able to show quantitative and qualitative test results.
Additionally depending from the limits setting, results are displayed in green (PASS), yellow	(ALERT) and red (FAIL) colours
Quantitative result indications ; Active, Fast, Passive and Mouth alcohol modes
:

---
Alcohol level indication
Qualitative result indications for scanning and mouth alcohol modes

---
Alcohol level indication
The following test results are displayed when “P/A/F” – unit is selected

---
Navigation of the unit

---
Measurement modes
Active mode (AM)
to determine the precise alcohol level in the expired air; it is ensured that only deep lung air is	delivered to the sensor
Fast mode (FM)
this mode is intended for persons suffering a COPD (e.g. Asthma) or reduced lung capacity,	which might not be able to provide the required minimum volume; by default 1.2L
Passive mode (PM)
measurement is triggered manually by the operator of AlcoTrue P
Scan mode (SM)
sample uptake is manually activated by the operator of AlcoTrue P; after the manual	activation 10 samples are performed
Mouth alcohol mode (MA)
mouth alcohol mode consists of two tests with a determined waiting period between the first	and second measurement

---
Memory
Result indications
Test No (9999 values)
Date when test was performed
Measured alcohol concentration
Type of Alcohol concentration
(BAC blood alcohol concentration
or BrAC breath alcohol concentration)
Time when test was performed
Type of measurement mode
Calibration status
Test location*
Print option*
1
2
1
4
3
2
5
4
3
6
5
7
6
8
7
9
8
*Features available upon request
9

---
Printing of results
AlcoTrue P printer
The results can be printed using bluepoint MEDICAL AlcoPrint
Printing option of AlcoTrue P can be activated/deactivated only by	authorised service	personnel
A
B
It is possible to print a test result immediately after
a measurement is complete (left screen) or from the
“Results” menu (right screen)
Indicates printing process

---
AlcoTrue P - Service Levels
Service level 1
  - Access to the service menu
  - Service menu and service menu setting
  - Alcohol calibration
Service level 2
  - Service level 1
  - Exchange hardware components ( housing parts, fuel cell)
  - Data export
Service level 3
  - Service level 2
  - Exchange all hardware components
  - Access factory setting
  - SW update

---
AlcoTrue P - Service equipment
Service level 1
  - Wet gas calibrator ( 34°C )
  - Pump
  - Flowmeter
  - Calibration solution ( 1-2 L depends from the calibrator type)

---
AlcoTrue P - Service equipment
Service level 2
(service level 1 plus hardware components exchange)
  - Screw driver set
  - Electrical screw driver with Nm
  - Soldering station

---
AlcoTrue P - Service equipment
Service level 3
(service level 2 plus all hardware components exchange)
  - Factory access
  - Factory setting
  - Factory software applications
  - SW update

---
AlcoTrue P - Service levels 1 and 2
1. AlcoTrueP – Service menu access
  Step 1: Select “Service” from user menu
  Step 2: Edit 4 digits PIN ( 0 0 0 0)

---
AlcoTrueP – Service menu points
- Date/Time - set device date and time
- Calibration – Alc. Calibration
- Language – set device language
- Unit – set device unit
- Limits – set country specific alc. limits
- Sensor exchange – edit fuel cell S/N
- Report configuration – select additionally printout report points
- Data memory – delete all saved results
- PIN change – change service PIN
- Printer – enable/disable print option
- Last results – set No. of last shown results
- Modes – enable/disable available test modes

---
AlcoTrueP – Calibration
alcohol calibration
  - Wet gas ( default)
  - Dry gas (upon customer request)
Flow calibration
  - Necessary only if pressure sensor exchange ( service level 3)

---
AlcoTrueP – Wet gas calibration procedure
Step 1: Turn on device
Step 2: Go into service menu
Step 3: Select Calibration – Wet gas
Step 4: Connect device to the Calibrator
Step 5: Start the calibrator
Step 6: Stop the calibrator when the signal ton stop
Step 7: Set the calibration point to the calibration solution
(Attention: ensure that device measurement unit and calibration concentration unit are same)
Step 8: Confirm the change
Step 9: Calibration successful / Calibration failure
Step 10: Escape from service menu
Note : Allowed tolerance +/- 3% from measured value by BAC 1,00‰ ( BrAC 0,476mg/L)

---
AlcoTrue P – service Level 1 and 2 ( Exchange Hardware components)
AlcoTrue P - housing open
Step 1: Remove screw secure
Step 2: Remove the screws

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)
AlcoTrue P - inside look

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)
AlcoTrue P - main components

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service Level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service Level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service Level 1 and 2 ( Exchange Hardware components)

---
AlcoTrue P – service Level 1 and 2 ( Exchange Hardware components)

---
Thank your
for you attention!