	Microstim DB3 Output Specifications

	[image: image1.jpg]




	Train of Four

Four pulses at 2 Hz.

Pulse width:                                                   200 μs

Pulse interval:                                               500 ms

Number of pulses:                                                  4

	[image: image2.jpg]TRAIN OF FOUR

[~ 500 ms—





	Double Burst Stimulation

Two bursts of pulses at 50 Hz, 0.75 seconds apart. 

First burst, 3 pulses followed by second burst of 2 pulses e.g. ratio of 3:2.

Pulse width:                                                   200 μs

Pulse period:                                                  20 ms

Duration between pulses:                             750 ms


	[image: image3.jpg]DOUBLE BURST STIMULATION

=20 ms-| —— 750 ms ——|





	Post Tetanic Count

5 seconds at 50 Hz (Tetanic), then a pause of 
3 seconds; followed by continuous 1 Hz stimuli, duration dependent upon switch depression.

Pulse width:                                                   200 μs

Pulse period, Tetanic:                                    20 ms
Pulse period, after recovery period:                  1 s


	[image: image4.jpg]POST TETANIC COUNT

J 1|11

3s }—1 s-{
Tetanlc Recovery Single Twitches
50 Hz 1 Hz, Continuous






	1 Hz

Continuous 1 Hz stimuli, duration dependent upon switch depression.

Pulse width:                                                  200 μs

Pulse interval:                                           1 second

Number of pulses:                                 Continuous


	[image: image5.jpg]




	Maximum output voltage: >90 V p-p

	Battery LED voltages

Red:      <6 V

Yellow:  >6 V, <8 V

Green:   >8 V


Part Number: 2590011
v1.0
9th September 2015


