A Register of measuring equipment will be prepared and maintained  in Intrastats replacing (Form QC14


This Register will contain a list of all the measuring equipment in the Company, stating the Serial No, if known location and frequency of calibration.


Most equipment is Indication Only and is only calibrated before use when the results are critical�A Calibration Record Card (Form QC15) will be kept for equipment calibrated in-house, which  will  list the following:- 


Serial Number of instrument


Type of Instrument and make


Test (Range)Standard used if any


Tester Printed name and signature and Date


The result of each calibration will be recorded on the Record Card, indicating acceptance or not and will be signed and dated, with date for next test.


Each piece of equipment will be labelled, either on the equipment itself or on its protective case,showing the calibration status of that equipment, colour coded as follows


Green                Externally Calibrated


Yellow               In-house calibrated


Red                    Indication only not calibrated.


BARCODE �Barcode Tracking Sticker will be applied to all Equipment


Barcode Tracking Sticker will be applied to locations


Equipment not calibrated by Viamed will be sent to a sub-contract source capable of issuing anapproved calibration certificate.


The intervals between calibration will be set by the Managing Director and will be related to the particular piece of equipment and its frequency of use.  Normally this will be every two (2) years


Subsequent changes to the calibration interval can only be made by the Managing Director or a designated person


 These will relate to the deviations found from previous calibrations. 


Additional calibration checks will be made if:


The equipment has been dropped or damaged,


The user has reason to suspect the equipment’s accuracy,


The user needs accuracy beyond indication


Any equipment which cannot be re-calibrated will be withdrawn for repair, and if irreparable will be replaced or scrapped. 


If any item of equipment is found to be out of calibration the Managing Director will assess the


	effect that this could have had on products previously checked by that equipment.  He will decide


	if the customer shall be informed that he has received suspect products or if a recall is necessary.


              

















Calibration Register


Calibration Register is now Tracked from Within Intrastats,


Triggered From Management Meeting OR Manually Triggered Via Intrastats -> ISO -> Calibration Index. (Fig 1.0)





�


Fig 1.0





Initial List shows Equipment where Next Due date is Past and Equipment Location (Fig.2). To see a complete list of all Equpiment in the Calibration File and Locations Click ‘List all Calibration File’.





�


Fig 2.0


To request an Employee sorts out a calibration Click ‘Request Calibration’ – This will open an issue and prefill the Issue text (fig.3)


�


Fig 3.0


When Calibration has been Completed Click ‘Calibrate”    Where???





�





Fig 4


Fill in Fields to Log the Calibration and any Certificate numbers.


COMPANY OPERATING PROCEDURES�
�
Calibration�
�
Created:�
17/May 1995�
VM3/COP/11�
Issue �
3�
�
Revised:�
�date \@ "dd MMMM yyyy" \* MERGEFORMAT�12 October 2006��
Last printed � PRINTDATE �15/06/2006 13:05��
Page � PAGE �3� of � NUMPAGES �3��
�






� FILENAME �Modified Audit 6 calibration.doc�                      � DATE \@ "dd/MM/yyyy" �12/10/2006�                      Page � PAGE �3� of � NUMPAGES �3�











