## Oximeters, Pu
Manufacturer and Model Name |  |  | Sample Column | Viamed VM-2101 Finger Pulse Oximeter | Viamed VM-2103 Finger Pulse Oximeter | Viamed VM-2105 Silicone Finger Pulse Oximeter | VM-2160 Handheld Pulse Oximeter
WHERE MARKETED |  |  |  | UK, Europe, Internationally (excluding USA) | UK, Europe, Internationally (excluding USA) | UK, Europe, Internationally (excluding USA) | UK, Europe, Internationally (excluding USA)
FDA CLEARANCE |  |  | Yes/No/Submitted | No | No | No | No
CE MARK (MDD) |  |  |  | CE0086 | CE0086 | CE0086 | CE0086
STAND-ALONE/ MODULAR |  |  |  | Stand-alone | Stand-alon | Stand-alon | Stand-alone
MOUNTING OPTIONS |  |  |  | None | None | None | Rail mount, pole mount, medirail mount, universal clamp, wall plate
DISPLAYS |  |  | SpO2, pulse rate, pulse strength and/or signal, low battery | Oxygen saturation, pulse rate, phethysmogram, pulse indicator, bar graph, pulse signal quality, signa strength, battery status | Oxygen saturation, pulse rate, phethysmogram, pulse indicator, bar graph, pulse signal quality, signa strength, battery status | Oxygen saturation, pulse rate, phethysmogram, pulse indicator, bar graph, pulse signal quality, signa strength, battery status | Oxygen saturation, pulse rate, plethysmogram, bar graph, short and long term trends, signal qualtiy, pulse amplitude, battery status, alarm silenced, sensor detection, sensor disconnection
| Type |  | LCD, LED, others | OLED | OLED | OLED | TFT
SpO2 RANGE, % |  |  | 70-100 | 0-100 | 0-100 | 0-100 | 0-100
| Accuracy, % |  | 70-100 ±3 | ARMS=2.0%, +/- S.D. = 1.86% | ARMS=2.0%, +/- S.D. = 1.86% | Grey: 1.519 ARMS, +/- S.D. = 1.513% | +/- 2%
PULSE RATE, bpm |  |  | 30-240 | 20-300 | 20-300 | 20-300 | 20-300
| Accuracy |  | ±5 | +/- 1 digit (</ 100 bpm) +/- 1% (> 100 bpm) | +/- 1 digit (</ 100 bpm) +/- 1% (> 100 bpm) | +/- 1 digit (</ 100 bpm) +/- 1% (> 100 bpm) | +/- 1 digit (</ 100 bpm) +/- 1% (> 100 bpm)
PERFUSION INDEX AND/OR SIGNAL STRENGTH INDICATOR |  |  | Low perfusion technology | Low perfusion technology and signal strength indicator | Low perfusion technology and signal strength indicator | Low perfusion technology and signal strength indicator | Low perfusion technology and signal strength indicator
RESPONSE TIME, sec |  |  | Depends on application | NA | NA | NA | NA
START-UP TIME, sec or pulses
AVERAGING TIME, sec |  |  |  | 3 seconds | 3 seconds | 3 seconds | 3 seconds
ALARMS
| Audible and visual |  | High/low SpO2 and pulse rate, sensor off, sensor failure, low battery |  |  |  | High/low SpO2 and pulse rate, pulse amplitude
| Visual only |  |  | Pulse signal quality, signal strength, battery status, low battery | Pulse signal quality, signal strength, battery status, low battery | Pulse signal quality, signal strength, battery status, low battery | sensor failure, low battery, battery status, signal quality, sensor detection, sensor disconnection, sensor off, alarm silenced,
ALARM OVERRIDE
| Reactivation method |  | Automatic or manual | NA | NA | NA | Manual
| Volume control |  |  | NA | NA | NA | Yes
SELF-TEST MODE |  |  | Preferred
PROBE TYPES
| Disposable |  |  | NA | NA | NA | Available
| Reusable |  |  | NA | NA | NA | Available, silicone included as standard
| Patient range |  | Adult, pediatric, infant, neonate | Adult Paediatric | Adult Paediatric | Adult Paediatric | Adult Paediatric Infant Neonatal
| Cable length, m (ft) |  |  | NA | NA | NA | Type dependant, between 0.45m and 1.2m
DATA MANAGEMENT
| Data stored |  | SpO2, pulse rate | NA | NA | NA | SpO2, pulse rate
| Data storage, hr | [5 empty] | 480
| External output | [5 empty] | USB to PC
PRINTER/RECORDER |  |  | Optional |  |  |  | NA
MRI CONDITIONAL |  |  | Yes/No
LINE POWER, VAC |  |  | 100-200, 50/60 Hz |  |  |  | Battery powered
BATTERY TYPE (NUMBER) | [6 empty] | AA alkaline batteries
Life, hr
| Rechargeable |  | Yes, No |  |  |  | No
|  | Recharge time, hr |  |  |  |  | NA
| Low-battery notice |  | Yes, No |  |  |  | Yes
H x W x D, cm (in) | [6 empty] | 11.8 x 6 x 2.5
WEIGHT, g (oz) | [6 empty] | 160
PURCHASE INFORMATION
| List price | [5 empty] | 345
|  | Replacement probe |  |  |  |  | Silicone Adult £174.25
| Warranty | [5 empty] | 24 months
Delivery time, ARO
| Year first sold |  |  | 2008 | 2012 | 2009 | 2007
Fiscal year
GREEN FEATURES |  |  | Please list any environmetally friendly features of this device.
OTHER SPECIFICATIONS
UMDNS CODE(S) |  |  | 17148 | 17148
LAST UPDATED
Supplier Footnotes
Model Footnotes
Data Footnotes
Note on A7: The parameters displayed are an important consideration because the device's monitoring capabilities determine its usefulness in a facility.
Note on A9: The range of arterial hemoglobin (SpO2) that the unit is capable of measuring. Facilities should choose models that have adequate measuring capabilities for their intended area of use.
Note on A11: The pulse rate is the frequency at which the heart beats. (Note: physiologic monitors obtaining heart beats via electrical conductivity of the heart [e.g., ECGs] are generally termed heart rate. Pulse rate generally refers to the measurement of heart beats obtained by counting physical/mechanical contractions of the heart. For this reason pulse oximeters display pulse rates and not heart rates.)
Note on A13: Perfusion index, sometimes called perfusion indicator, is an indication of signal strength. It is usually displayed as a numeric value. (It could also be represented graphically). The perfusion index is given as a range, and the higher the reading the more confidence in the saturation measurement. Likewise, the lower end of the range is important, especially if the manufacturer can indicate the oximeter's perfusion index threshold level (e.g., how weak a signal can be and still get a valid saturation measurement). A manufacturer may provide a signal-strength indicator in addition to, or instead of, a perfusion index.
Note on A14: Time lapse between time of measurement and the update of the SpO2 display. In some units, the user can select how frequently the SpO2 readings will be updated.
Note on A15: Time elapsed between unit activation to first SpO2 reading being displayed.
Note on A16: Movement of the sensor, clinical or otherwise, may be detected and displayed by the oximeter as artifact and can cause inaccurate readings. The effects of interference caused by motion can be reduced, or even eliminated, by digital signal processing and by averaging the displayed SpO2 over several seconds.
Note on A34: Whether the pulse oximeter's design features (e.g., fiberoptic cable) allow its safe use during magnetic resonance imaging studies. Buyers should carefully review MR compatibility claims to ensure the pulse oximeter being selected can be used as intended within the MR environment if applicable.
Note on A50: If the model has any environment-friendly features, components, or specifications, they are listed here. This field can also contain manufacturer-specific information, such as green-certified production or recycling methods.
Note on A52: Universal Medical Device Nomenclature System™ (UMDNS™) is ECRI Institute’s nomenclature and computer coding system for medical devices. It is used in ECRI Institute’s databases and publications, as well as in thousands of healthcare institutions worldwide. The 5-digit code listed for each model in the chart refers to the nomenclature term that applies to that model; in some cases, a single chart will include devices that fall under different nomenclature codes (for example, an automated device versus a semiautomated device). For more information on UMDNS, see the About UMDNS page on the HPCS homepage.
Note on B10: Accuracy is based on manufacturer data (in some cases under ideal conditions) and may not reflect pulse oximeter performance during actual conditions.
Note on B22: In some units, the volume of the alarm and pulse tone is user adjustable. Other units have an internal setting that can be adjusted only by the hospital engineer.
Note on B27: Lists probes available for each patient type. Not all manufacturers have specified whether their probes can be used on neonates. Buyers should check with the manufacturer to determine whether probes have been tested for use with appropriate patient type. Both disposable and reusable, sensor probes are available for many oximeters.
Note on B37: The amount of time, in hours, that the unit battery will last before it must be replaced. Buyers should select a model that can provide adequate battery life for its intended use.
Note on B44: Indicates the list price or price range of the model. List prices do not reflect customization, group purchasing or other manufacturer agreements, or other discounts. ECRI Institute believes that although list prices enhance equipment planning projects, they should not be the sole consideration when making a purchase decision.