VM-2400
Multi parameter Monitor for SpO2, non-invasive Hemoglobin and Methemoglobin

VM-2400 User Manual

Version:
Released:

The information in this document is subject to change without notice.

Copyright © 2016 Viamed Ltd. All Rights Reserved.
This document contains proprietary information that is protected by copyright. All rights are reserved. No part of this document may be photocopied, reproduced or translated to another language without prior written consent of Viamed Ltd.

Contact information

Viamed Ltd.
15 Station Road
Cross Hills
Keighley
West Yorkshire BD20 7DT
United Kingdom

Tel:	+44 (0) 1535 634542
Fax:	+44 (0) 1535 635582
E-mail:	info@viamed.co.uk
www.viamed-online.com

Declaration of Conformity

CE 0482
This product meets the requirements of the 93/42/EWG guidelines for medical products.

Liability
Viamed Ltd. shall in no event be liable for any direct, indirect, special or consequential damages including without limitation damages for loss of business profits, loss of income, business interruption, loss of business information, loss of use or other related exposures, however caused, arising from the faulty or incorrect use of the product.
Viamed Ltd. guarantees that the product delivered has been thoroughly tested to ensure that it meets its published specifications.

Warranty
Viamed Ltd. warrants the products manufactured or distributed by them to be free from faulty materials and workmanship for a period of 24 months from date of original shipment to first end user except for disposable products or products which have a stated guarantee longer or shorter than 24 months. Viamed Ltd. will perform warranty service at its factory.
The obligations for Viamed Ltd. under this guarantee shall be limited to repair, or at Viamed Ltd. 's discretion, replacement of necessary parts or assemblies and shall not include shipping costs or other incidental damages.
Claims for damages during shipment must be filed promptly with the transportation company. All correspondence concerning the products must specify both the name of the product and its serial number as written on the label on the product.
Use of the equipment for other than its intended use, or if it has been repaired by anyone except Viamed Ltd. or a Viamed Ltd. authorized service centre, or altered or modified or used without following the instructions in the user manual, will void this guarantee.

Table of Contents

# Introduction
## Intended Use
The VM-2400 multi parameter monitor is intended for non-invasive spot check measurement of SpO2, non-invasive Hemoglobin (Hb), Hematocrit (Hct), Methemoglobin (MetHb) and Pulse Rate (PR) of adults and paediatrics.

VM-2400 and accessories are intended to be used by trained professionals only.

The VM-2400 may be used in clinical and non-clinical settings e.g. hospitals, hospital type facilities, operating suites, intensive care units, physician offices, ambulatory surgery centers and blood donation facilities.
## Contraindications / Hazards
There are no known contraindications for patient monitoring with the VM-2400, provided that the data obtained is evaluated with consideration given to the patient's clinical condition.
For further information regarding possible hazards during operation of the device - see section Warnings.

## Warnings
Warning:
Warnings are identified by the WARNING symbol shown above. Warnings alert the user to potential serious outcomes, such as death, injury, or adverse events to the patient or user.

Warning:
Do not make any clinical judgments based solely on the VM-2400. The monitor is intended only as an adjunct in patient assessment. It must be used in conjunction with clinical signs and symptoms. The interpretation of the measurement values should be done only by trained health care professionals.
Warning:
Explosion hazard - Do not use the monitoring system in the presence of flammable anaesthetics mixture with air, oxygen, or nitrous oxide.
Warning:
Explosion shock hazard - Do not open the VM-2400 device. Only trained experts may perform service procedures.
Warning:
Do not place containers with liquids on or near the VM-2400 - Liquids spilled on the instrument may cause it to perform inaccurately or fail.
Warning:
Routinely monitor the patient to make sure the VM-2400 is functioning and the sensor is correctly placed.
Warning:
Variation in hemoglobin measurements may be affected by physiological conditions, sample type and body positioning. As with most hemoglobin tests the VM-2400 results should be checked with the background on the patient’s conditions. Each result which contradicts with the patient’s clinical status should be repeated and/or complemented with additional test data.
Warning:
VM-2400 can show faulty measurements or can be damaged if it will be used outside the specification or environmental conditions.
Warning:
Certain physiological conditions, medical procedures, and external agents may interfere with the monitor‘s ability to detect and display accurate measurements. (Chapter General Information provides information on possible interferences)
Warning:
If patient hypoxemia is indicated, blood samples should be analysed by laboratory devices to completely understand the patient’s condition.
Warning:
For the measurement, the Hb monitor uses five wavelengths in the range of 600nm to 1400nm. Consider that these wavelengths might influence diagnostic parameters of other optical applications.
Warning:
If you are uncertain about the accuracy of any measurement, check the patient’s vital signs by alternate means; then make sure the VM-2400 is functioning correctly.
Warning:
The use of accessories, sensors, and cables other than those specified may result in increased emission and/or create invalid readings of the VM-2400.
Warning:
The VM-2400-Sensor is designed for use with the VM-2400. Verify the compatibility of the monitor, sensor and cable before use, otherwise patient injury can result.
Warning:
In high ambient light conditions it is required to shield the sensor application site with opaque material. Too much ambient light may result in inaccurate measurements.
Warning:
The VM-2400 is a prescription device to be operated only by trained personnel. The monitor is for attended monitoring only.
Warning:
No modification of VM-2400 and the accessories is allowed.

Warning:
As with all medical equipment, carefully route patient cabling to reduce the possibility of patient entanglement, strangulation, or injury to the patient
Warning:
Do not place the VM-2400 or accessories in any position that might cause it to fall on the patient. Do not lift the VM-2400 by the cable or sensor.
Warning:
The VM-2400 should not be used as a replacement or substitute for ECG based arrhythmia analysis.
Warning:
The VM-2400 is not defibrillator proof. Remove the sensors from the patient throughout defibrillation or while an electrosurgical unit is in use to avoid shock of the caregiver or patient.
Warning:
Disconnect the VM-2400 and sensor from the patient throughout computed tomography or magnetic resonance imaging (MRI) scanning. Induced current could potentially cause burns.
Warning:
EMI radiation interference such as LCD/plasma TVs, computer displays cause errors or incorrect measurements on the VM-2400.
Warning:
Do not use sensors or cables that appear to be damaged. Do not use a sensor when optical components are exposed. Do not use a device that appears damaged. Replace monitor immediately in cases of visible damages.
Warning:
Any misapplication of the sensor with excessive pressure for prolonged periods can induce pressure injury.
Warning:
If using the VM-2400 during full body irradiation, keep the sensor out of the radiation field. If the sensor is exposed to the radiation, the reading might be inaccurate or the instrument might read zero for the duration of the active irradiation period.
Warning:
Do not place the VM-2400 where the controls can be changed by the patient.
Warning:
Do not place the VM-2400 on electrical equipment that may affect the instrument, preventing it from working properly.
Warning:
Do not expose the VM-2400 to excessive moisture such as direct exposure to rain. Excessive moisture can cause the instrument to perform inaccurately or fail.
Warning:
Do not damage batteries or apply pressure. Do not hit or throw the batteries.
Warning:
Do not immerse the battery in water or seawater. Store it in a cool and dry environment if not used.

Warning:
Do not discard, store or use the battery near a heat source (e.g. a fire or heater).

Warning:
Only charge the Li-Poly battery CT-2500 while inserted in the VM-2400 using the provided power supply FW 7660M/06.

Warning:
Do not connect the positive and negative terminal with metal objects such as wire and do not transport or store the battery together with metal objects such as necklaces or hairpins as this may short-circuit the battery.

Warning:
Do not strike, throw or trample the battery or pierce it with a nail or other sharp object.

Warning:
Only use the Li-Poly battery Model No. CT-2500 provided with the monitor. Contact the manufacturer for replacements.

Warning:
Only use the power supply Model No. FW 7660M/06 provided with the monitor. The use of a power supply other than this may result in hazardous situation and effect patient’s safety.

Warning:
Do not use rechargeable AA sized batteries instead of alkaline AA sized batteries to operate the device, as this may affect the function of the device.

Warning:
Batteries may leak or explode if used or disposed of improperly.

Warning:
Explosion hazard — Do not use AA rechargeable Lithium-ion batteries.
Use only Alkaline (1.5 Volt, Type AA LR6) or rechargeable NiMh AA (1.2 Volt, Type AA HR6, 1200 mAh)

## Symbol Description
	Accompanied with “Warning: Supplementary text.” within this document. Warnings indicate potential harmful conditions that may lead to injury or death.
	Accompanied with “Caution: Supplementary text.” within this document. Cautions indicate conditions that may lead to damage to or malfunction of the device.
Note:	Denoted as “Note: Supplementary text.” within this document. Notes inform the user to relevant facts and conditions in connection with the device.
	Consult User Manual for detailed operating information.
	Consult accompanying documents for important safety-related information
	Manufacturer
	Type BF applied parts
	Catalogue number
	Serial number
	Temperature limitation
	Indicates that the transport package has to be kept out of heated areas.
	Humidity limitation
	Atmospheric pressure limitation
	Fragile, handle with care
	Keep dry
	Do not dispose in the consumer waste. Electrical and electronic equipment shall be collected and recycled in accordance with (Directive 2002/96/EC)
IP31	IP Code (International Protection Rating) according to IEC 60529 IP3X: protection against tools, wires or similar objects with Ø > 2.5 mm, protection against solid foreign bodies with Ø > 2.5 mm / IPX1: Protection against vertically falling water drops.
	No parameter alarms

## Terms and Definitions
VM-2400	Handheld Hb monitor
VM-2400 - Counter Key	Measurement payment system
VM-2400 – Sensor	Hb sensor
RHb	Deoxygenated haemoglobin
Hb	Haemoglobin
HCT	Hematocrit
MetHb	Methemoglobin
SpO2	Functional arterial oxygen saturation
PR	Pulse rate
PI	Perfusion Index
IR	Infrared
Li-Poly battery	Lithium-poly rechargeable battery
MDD	Medical Device Directive
MRI	Magnetic Resonance Imaging
RH	Relative humidity
USB	Universal Serial Bus

# Theory of operation
## Hb and MetHb measurement
The non-invasive measurement of SpO2, hemoglobin and its derivative MetHb with the VM-2400 is based on the fact that different hemoglobin derivatives absorb light at specific wavelengths. The extinction coefficients for methemoglobin (MetHb), oxyhemoglobin and reduced haemoglobin are shown in the figure below.

Figure : Extinctions Coefficient

The basic principal of VM-2400 is a five wavelength measurement in the range of 600nm to 1400nm. Measurement of Hemoglobin and MetHb is based on same principles as measurement of functional oxygen saturation. For all wavelengths the change in absorption is continuously measured in transmission mode. In this way the pulsatile signal due to the arterial blood alone is extracted, excluding the offset due to absorption by venous blood, skin, bone, muscle and fat. Using the multi wavelength calibration equation the Hb and MetHb is calculated.

## SpO2 measurement
Measurement of functional arterial oxygen saturation (SpO2) is based on the fact that the amount of absorbed infrared light and red light is different for oxygenated haemoglobin (HbO2) and deoxygenated haemoglobin (RHb).
The amount of red and infrared light received is compared and the percent of haemoglobin molecules bound with oxygen molecules calculated:

These measurement values are continuously displayed as a waveform (plethysmogram) and also used to determine the pulse rate.
Certain physiological conditions, medical procedures, and external agents may interfere with the monitor’s ability to detect and display accurate measurements. See chapter 7.1 for detailed information.
In ambient temperature 35°C the maximum skin surface temperature at the pulse oximeter probe has been measured at less than 41°C according to the ISO 80601-2-61 Annex BB.

# Product overview
The VM-2400 is used together with the VM-2400 - Sensor and the VM-2400 - Counter Key. Power is delivered via the provided power supply (FW 7660M/06), a rechargeable Li-Poly battery (CT2500) or 4 x alkaline batteries (1.5 Volt, Type AA LR6).

Key features of VM-2400:
- Simple and intuitive handling
- Fast measurement (approx. 40s)
- Detection of unphysiological MetHb
- Pain Free
- Long battery life (up to 300 measurements)
- Storage of up to 12.000 measurement
- Reliable
- No bio-contaminated waste
Figure : VM-2400-Sensor	Figure : VM-2400 - Counter Key

VM-2400 enables simple and fast measurements.
Based on many years of experience gained in pulse oximetry a new generation of non-invasive handheld diagnostic tools was developed. Compared to laboratory diagnostics, the new VM-2400 offers many advantages, including increased efficiency and reduced operating costs. The simple and intuitive operation enables fast measurements, without the use of needles or fleams and does not require any expert knowledge. With this non-invasive procedure the patient will experience a pain-free and stress-free blood analysis, without the risk of cross-contamination. In addition there are environmental benefits from the reduction in contaminated bio-medical waste.
High quality and comfortable sensors that incorporate the very latest technology are used:
Due to the design, the VM-2400 - sensor is robust and can be comfortably used with a wide range of finger sizes. Reliable elimination of even the most problematic pathogens using highly effective disinfectants. The VM-2400 sensor is manufactured from premium material and their design allows for effective high-level disinfection. This reduces the risk of nosocomial infection associated with surface-borne pathogenic microorganisms.

Methemoglobin:
MetHb is an abnormal hemoglobin which is not able to bind oxygen and reduces the oxygen uptake capacity. High and non-physiological MetHb values reduce the blood quality significantly and the donated blood should not be used for blood transfusions. The formation of MetHb can be caused by local anesthesia, several medications or drugs. Excluding blood donors with high MetHb concentrations can raise the quality of blood preservations.

# Controls – Symbols – Display Modes
## Controls and Unser Interfaces
Front View	Side view	Top view	Rear view

No.	Symbol	Features/Button	Function
		Arrow Buttons / (up/down)	Multifunction button used for: / scrolling through menu items and / increasing/ decreasing parameters. / from monitoring display modes: can be used as shortcuts to volume control
		Result history	Display the last 10 measurement results
		Gender Button (male)	Male selection
		Menu	Menu selection
		ENTER Button, On/Off	Confirms selection / To turn on the device: press and hold power button briefly. / To turn off the device: press and hold power button for approx. 3 seconds.
		Gender Button (female)	Female selection
		USB	USB 2.0 interface
		Counter Key Port	Port for VM-2400 - Counter Key / Warning: / Only connect VIAMED VM-2400 - Counter Key.
		Sensor Port	Port for VM-2400-Sensor / Warning: / Only connect VIAMED VM-2400-Sensor
		Power input	Port to connect the external power supply (100-240V AC / 50-60Hz, Model No. FW 7660M/06)

## Display Modes and Display Data

Start Screen 1	Start Screen 2- female selected – finger inserted	Start Screen 3- male selected – finger inserted

Analysing Screen	Result Screen	Result history (last ten results)

- SpO2	(functional saturation)
- PR	(pulse rate in beats per minute)
- PI	(perfusion index)
- Hb	(non-invasive Hemoglobin in g/dL, g/L and mmol)
- Hct	(Hematocrit[1])
  [1] Calculated from Hb value (Hb [g/dl] * 2.94) | J.D Bauer, P.G Ackermann, G. Tori, Clinical laboratory methods. The C.V. Mosby Company, Saint Luis 1974, S. 156
- MetHb	(Methemoglobin)
Gender selection
Before measurement select the patient’s gender at the keypad.
## Symbols and Indicators

No.	Symbols/Indicators	Definition
	/	Battery level indicator. The three segments represent the battery charge level. The symbol flashes yellow and audible indication when the battery capacity is low. The symbol flashes red when the battery capacity is near to empty, replace AA batteries or charge Li-Poly battery. / Plug symbol if AC power is connected and no batteries inserted.
		Current time, displayed in 12h or 24h format
		Available Spot Check Measurements (VM-2400 - Counter Key)
		Indicates the dynamic pulse amplitude and rate. As the detected pulse becomes stronger, more bars light with each pulse. The reverse is true for weak pulses. The colour of the bar graph is an indicator for signal quality. / Green: good signal quality, very accurate measurement / Yellow: average signal quality, measurement may be inaccurate / Red: poor signal quality, unreliable measurement
		Plethysmogram, displays the normalized pulse waveform. / The reading is automatically adjusted to the pulse strength; therefore, a waveform with strong amplitude should be visible at all times.
		Symbol of selected gender
		Memory full symbol. The oldest data is overwritten, after upon confirmation by the user.
		Analysing bar indicates status of spot check analysis. Total analysing duration is 40 sec.

# Preparation for Use
Caution:
The use of rechargeable instead of alkali AA batteries may cause a shorter operating time of the device.
## Power Supply
Warning: / Batteries may leak or explode if used or disposed of improperly. / The external power supply (100-240V AC / 50-60Hz, Model No. FW 7660M/06) is used for continuous operation of the monitor and to charge the Li-Poly battery. / To operate by mains, connect the power supply cable into the power input socket located at the back of the device. / Ensure that the correct power supply plug is connected to the power supply. It can be exchanged by pressing the release button (2.1) on the power supply. As standard the device is supplied with a European plug. Additional plugs are available upon request. / Connect the power supply to an AC outlet.	Figure : Connect the power supply

## Battery Installation
Warning: / Only charge the Li-Poly battery CT-2500 while inserted in the VM-2400 monitor using the provided power supply FW 7660M/06. / Slide down the cover of the battery compartment on the rear panel of the device. / Insert four alkaline batteries (1.5V, AA), ensuring the correct orientation in accordance with the polarity markings. / Alternatively, insert the rechargeable Li-Poly battery (Model No. CT-2500), orientated according to the guiding grooves. / Slide the battery-compartment cover back into its initial position to close. / Caution: / If the VM-2400 is not used for a longer period, the AA batteries must be removed and stored, to prevent damage of device by leakage of batteries. / Caution: / Check charge status of Li-Poly batteries on regulatory base, if stored outside the device.	Figure : Insert a Li-Poly battery or 4 x AA alkali batteries

## Connecting the Sensor and VM-2400 - Counter Key
Insert the sensor cable into the sensor port located on the top right edge of the device, ensuring correct orientation of the sensor connector and the port. / Insert the VM-2400 - Counter Key into the VM-2400 - Counter Key port located on the top left edge of the device, ensuring correct orientation of the VM-2400 - Counter Key and the port. / Caution: / Visual check: Before beginning operation, ensure that the device and sensor are not damaged.	Figure : Connect VM-2400 - Counter Key and MW-Sensor

Caution:
Use only VM-2400 - Counter Key. Never connect anything else than the Viamed VM-2400 - Counter Key to the VM-2400 - Counter Key port

## Switching on the Device
Warning:
During the power-on self-test of the device a single loud tone can be heard. If not, it is absolutely necessary to ensure that the speaker is not obstructed or defective!

Press and hold button briefly until an opening ‘welcome screen’ appears. After the power on self-test is successfully completed the device is ready for monitoring.

## Connection the Sensor
Attach the sensor to the middle finger of the not dominate hand. Ensure that the finger attached correctly and the sensor is firmly plugged in.
Figure : VM-2400 with connected sensor	Figure : Right position of the finger within the sensor

Caution:
The finger may not stick through the sensor.
Caution:
Do not press the palm of the hand on the table. The hand should remain as relaxed and motionless on the table.
## Start Measurement
Caution:
It is recommended to perform the measurement while the patient is in a seated position.
Caution:
During the measurement the patient may not speak or move his arm, hand or fingers as this can disturb the measurement.
Caution:
Start measurement when the bar graph turns green/ yellow. In case the bar graph stays red see section 7.4 in user manual for options to increase signal quality.

Note: It is necessary to have remaining measurements.
- Press or button to select gender of the patient.
- Press the button to start the measurement.
- Measurement takes 40 seconds.
- After a successful measurement the results will be displayed.

Note: At each successful Hb measurement the Counter Key value is decremented by one.
Caution:
An arterial blood specimen for laboratory CO-Oximetry analysis may be considered to verify the oxygen saturation and hemoglobin values.
## Termination of Measurement
The measurement will be automatically cancelled in following cases: / Finger out / VM-2400 - Counter Key disconnected / VM-2400 - Sensor disconnected / Bad signal quality (Error Screen)	Figure : Measurement canceled!

## Test Results
After a successful Spot Check measurement the result screen will be displayed. / Press button to perform the next measurement. / Caution: / At each successful Hb measurement the Counter Key value is decremented by one.

Press button to view the results of the last ten measurements.

## Switching off the Device
Press and hold the button for approx. 3 seconds to switch off the device. The VM-2400 device will also power off automatically after 15 minutes when not in use.
# Screen Contents – Menu Structure
## Main Menu
All important and frequently used settings are accessible through the main menu, which can be opened by pressing the button.

Navigation the Menu
Use the buttons to scroll through menu items. The currently selected menu item is highlighted by a coloured frame. Press the button to confirm your selection.

Entering Data
In some submenus it is possible to adjust a certain parameter. In this case the parameter can be increased or decreased using the buttons. The value will increase or decrease more quickly when the respective button is held down. Press the button to confirm the new value.

Back
To go back one step in the menu select “Back” or press button.

Exiting Menu and Returning to Start Screen
Press button to return immediately to the start screen. If no button has been pressed for more than 30 seconds, the device will automatically return to the monitoring screen.

### Data Management
In the Data Management menu the last 12000 results can be considered.

Stored Data
The VM-2400 device can store up to 12000 results. A new result is stored automatically after each measurement. Also the Patient ID will be stored if it was set before.
The device warns the user when the memory is almost full by displaying the symbol, if the maximum of 12000 is reached. After this maximum has been reached the oldest result is overwritten upon confirmation by the user. Stored results can be retrieved and erased under the menu item “Data Management”.
### Setup

Use the buttons to select an item, confirm with the button

#### Mode
In this menu the measurement mode can be adjusted. The standard VM-2400 Model is limited to Hb Spot Check Mode.
#### Limits
The limits enable a coloured presentation of the results. Values in the limits are displayed green values outside the set limits are displayed in yellow. / This feature is useful for fast result interpretation in blood donation centres. / Adjust the limits for hemoglobin (Hb) for each gender separately by using the buttons. Confirm setting by pressing button. / It is also possible to set a limit for the displayed methemoglobin (MetHb) result. / To switch of the limits press buttons until „--" are displayed.

#### Unit

Adjust the unit of the measured hemoglobin (Hb) using the buttons. Confirm setting by pressing button.

#### Volume

Adjust the volume using the buttons. Confirm new setting by pressing the button.

#### Date-Time

Adjust the date and the time using the buttons. Confirm new setting by pressing the button. Date and time are not erased when the batteries are temporarily removed.

#### Language

Adjust the language by using the buttons. Confirm new setting by pressing the button. All messages and menus will be displayed in the selected language.

#### Service

Service: / The Service submenu is protected by a PIN code; only authorised service personnel can access this menu.

### Patient ID
The VM-2400 features the possibility of saving a 12 digit patient ID together with every created data file.

Note: After each measurement the Patient ID is set to zero.

# Error Messages – Problems – Corrective Actions
## General Information

Warning:
Physiological conditions, medical procedures, or external agents that may degrade measurement performance or affect the accuracy of the measurement include the following:

  - Incorrect applications of the sensor
  - Placement of the sensor on an extremity with a blood pressure cuff, arterial catheter, or intravascular line
  - Externally applied colouring agents such as nail polish or artificial nails,
  - Intravascular dyes,
  - Placement of the sensor on an extremity with blood flow restrictors (arterial catheters, blood pressure cuffs, infusing lines, etc.),
  - Failure to shield the sensor site in high ambient light conditions,
  - Excessive patient activity,
  - Low perfusion,
  - Venous pulsations,
  - Anemia or low hemoglobin concentrations,
  - Cardiac dysrhythmia like extrasystole or atrial / ventricular fibrillation
  - Dysfunctional haemoglobin, e.g. caused by a carbon monoxide intoxication
  - Electromagnetic interference,
  - Electrosurgical interference.

## Unsuccessful measurement
If one vital parameter could not be measured the result screen shows dashes instead. / Note: no Hb is displayed if SpO2 < 90 % or MetHb > 5 % as the measurement accuracy is not given within this range. / Caution: / If no Hb value could be measured the measurement counter is not decremented.

## Error Messages – Causes
“Connect valid Counter Key!
No valid VM-2400® - Counter Key is connected.
- Check VM-2400® - Counter Key connection. Replace the VM-2400® - Counter Key if required

“Counter Key expired!”
The VM-2400® - Counter Key is expired, no more measurements are available.
- Replace VM-2400® - Counter Key with a new one

“No sensor!”
The sensor is not connected properly to the device.
– Check sensor connection

“Sensor off!”
The sensor has been removed from the monitoring site.
– Check that the sensor is properly attached to the patient.

“Battery low!” battery symbol flashing red with audible indication
The battery is almost completely discharged.
– Replace batteries immediately

“Sensor fault!”
The connected sensor is either defective or not compatible with the device
– Check sensor

“Device defective!”
Fatal device error, e.g. if the device screws are opened
– Contact qualified service personnel or your local sales representative

“Measurement cancelled! Bad signal quality”
Poor-quality pulse signal, for example as a result of cold hands (low perfusion) or movement of the sensor during measurement.
– Retry measurement and prevent movement or preheat hand 2 minutes with the InfraCare® IR-A radiator, see section Bad signal quality.

“Data memory full!!!”
– The limit of 12000 saved results was achieved. Either the oldest results are overwritten or the memory can be erased in the Data Management, see section Data Management.

## Failure – Cause – Correction Action
Problem: There is no response to the Power button.
Cause – Corrective Action: Ensure that the Power button is fully depressed. The batteries may be missing, discharged, or oriented incorrectly. Install new batteries.

Problem: No pulse signal found or the pulse signal cannot be found anymore
Cause – Corrective Action: Check the patient. Check if sensor is applied correctly. Check sensor and extension cable connections. Increase perfusion, see section Bad signal quality. Test the sensor on another subject. Try another sensor or extension cable.

Perfusion may be too low for the monitor to track the pulse. Check the patient. Preheat the hand with the InfraCare® IR-A radiator, see section Bad signal quality. Test the monitor on yourself. Change the sensor site. Try another sensor.

Interference due to patient activity may be preventing the monitor from tracking the pulse. Keep the patient still, if possible. Verify that the sensor is securely applied and replace it if necessary. Change the sensor site.

The sensor may be too tight, there may be interference due to ambient light, or the sensor may be on an extremity with a blood pressure cuff, arterial catheter, or intravascular line. Reposition sensor, as necessary.

Electromagnetic interference may be preventing the monitor from tracking the pulse. Remove the source of interference.

## Bad signal quality
Low perfusion at the measurement finger results in bad signal quality which affects measurement performance. This is often the case if the hands are cold like typically the case in winter months.
The colour of the bar graph indicates signal quality already before starting the measurement.

Perform the steps below to increase the signal quality if the bar graph colour is yellow or red:
- Hold hands under warm flowing water for some minutes
- Use a squeezable anti-stress ball for approx. 1-2 minutes
- Heat the hand for 2 minutes with an IR-A radiator (e.g. Philips InfraCare® IR-A radiator)

If the signal quality does not allow accurate measurement results the measurement is interrupted and a “Bad signal quality” error screen is displayed.

## Other Problems
EMI (Electromagnetic Interference)
This device has been tested and found to comply with the limits for medical devices according to EN 60601-1-2, EN 80601-2-61 and the Medical Device Directive 93/42/EEC. These limits are designed to provide reasonable protection against harmful interference in a typical medical installation.

Due to the proliferation of radio-frequency transmitting equipment and other sources of electrical noise in healthcare environments it is possible that high levels of such interference due to close proximity or strength of a source may result in disruption of performance of this device. Examples of noise sources in healthcare environments that could cause electromagnetic interference include:

- Electrosurgical units
- Cellular phones
- Mobile two-way radios
- Electrical appliances
- High-definition televisions (HDTVs)

The pulse can be obscured by electromagnetic interference. During such interference measurements may seem inappropriate or the monitor may not seem to operate correctly.

Disruption may be evidenced by erratic readings, cessation of operation or other incorrect functioning. If this occurs, the operating environment should be surveyed to determine the source of disruption and the following actions taken to eliminate the source:

- Turn equipment in the vicinity off and on to isolate the offending equipment.
- Reposition or relocate the interfering equipment.
- Increase the distance between the interfering equipment and this equipment.

The monitor generates uses and radiates radio frequency energy. Failure to follow these instructions may cause harmful interference with other devices in the vicinity.

# Maintenance – Cleaning – Testing
## Maintenance
Warning:
Before maintenance, cleaning or disinfection ensure the monitor is turn off and not in use.

Caution:
There are no user-serviceable parts inside the VM-2400. The cover should only be removed by qualified service personnel.

Caution:
Follow local, state and national governing ordinances and recycling instructions regarding disposal or recycling of the device and device components, including batteries.

The monitor requires no calibration. The VM-2400 sensor requires calibration each year. Contact qualified service personnel or your local sales representative for calibration or if service is required. The contact address and phone numbers are listed on the last page.

Caution:
A functional tester cannot be used to assess the accuracy of a pulse oximeter monitor or sensor.

Test of the measurement accuracy
The only reliable method of testing the measurement accuracy of a pulse oximeter is the clinical validation of the measurement data, indicated by the system monitor with sensor on the basis of a blood gas analysis. During extensive clinical studies, the monitor combined with the approved sensors evidenced the accuracy required.

## Cleaning
Warning:
Before cleaning the monitor, always turn it off and disconnect the power supply.

Use a soft cloth dampened with either a commercial, nonabrasive cleaner, or a solution of 70% alcohol in water. Lightly wipe the surface of the monitor.

Note: Clean the VM-2400 sensor before use on the patient. Clean the VM-2400 once per week or more frequently if the handled by multiple users.

Caution:
Please clean contaminated and/or dirty equipment before returning, following the cleaning procedure described below.

Caution:
Do not use caustic or abrasive cleaning agents on the device or the sensors.

Caution:
Do not immerse the VM-2400 in any liquid. Do not spray, pour, or spill any liquid on the VM-2400, its accessories, connectors, switches, or openings in the enclosure as this may damage the monitor.
## Disinfection
Use a soft cloth saturated with a solution of 10% chlorine bleach in tap water.

Caution:
Do not sterilize the device with superheated steam or with hot air!

Caution:
Do not autoclave or immerse the device or sensors in liquid. Do not expose the device or components to excessive moisture or liquids.
## Order new VM-2400® – Counter Key
Caution:
The VM-2400 requires connection of a valid VM-2400 – Counter Key to perform measurements.

Caution:
Use only VM-2400® – Counter Key.

Each successful measurement the VM-2400® - Counter Key value is decremented by one. If no more measurements are available a message will be displayed. In this case please order a new valid VM-2400® - Counter Key. Counter Key with various numbers of measurements are available.

# Technical Specifications
Measurement Range:
SpO2:	0 – 100 %
PR:	30 – 240 bpm
Hb:	0 – 25 g/dl
Hct:	35 – 53 %[2]
[2] Calculated from Hb value (Hb [g/dl] * 2.94) | J.D Bauer, P.G Ackermann, G. Tori, Clinical laboratory methods. The C.V. Mosby Company, Saint Luis 1974, S. 156
MetHb:	0 – 10 %
Pi:	0.5 – 20 %

Accuracy:
SpO2:	60 – 100 % ± 2 Arms3
PR:	30 – 240 ± 2 bpm
Hb:	8 – 20 g/dl ± 1 Arms[3],[4],[5]
[3] Arms: ± 1 Arms represents approx. 68 % of measurements.
[4] Improved perfusion by IR-A radiator
[5] Valid if SpO2 > 89 % and MetHb < 6 %
Hct:	35 – 53 % 2.9 Arms2,3,4,5
MetHb:	0 – 10 % ± 1 Arms3

Arms referenced to SaO2 ranges:

100 – 60 %	100 – 90 %	90 – 80 %	80 – 70 %	70 – 60 %
1,63	1,31	1,91	1,56	1,67

Display:
- TFT colour graphic display, 262,000 colours, 240 x 320 pixels
- Data displayed: SpO2, non-invasive Hemoglobin (Hb), Hematocrit (Hct), Methemoglobin (MetHb), Pulse Rate (PR), plethysmogram, bar graph
- Indicators: Signal strength and signal quality, pulse amplitude, battery status, storage status, time, sensor detection, sensor disconnection

VM-2400® Sensor:

Maximum LED Power:	< 4 mW
Peak Wavelength:	600nm to 1400nm (5 selected wavelength)
Red Wavelength range:	600 nm to 700 nm
Infrared wavelength range:	800 nm to 1400 nm

Environmental Conditions:
Operating conditions:
0 to 40°C; 15 to 95% RH (non-condensing); 800 to 1200 hPa (excl. Li-Poly battery[6])
[6] With Li-Poly battery the conditions are reduced to: operating while charging at 0 – 40°C and 860 – 1060 hPa, storage (1 month) at -20 - 60°C and 860 – 1060 hPa

Storage conditions:
-30 to 70°C; 10 to 95% RH (non-condensing); 800 to 1200 hPa (excl. Li-Poly battery6)

Note: Should the device experience condensation it should be stored for more than 24 hours in an environment with relative moisture content below 95%RH (non-condensing).

Power supply:
Power may be supplied by battery, rechargeable battery or by AC power supply

4 x Batteries:
Alkaline battery (1.5Volt, Type AA / LR6 / AM3 / MN1500 / Mignon) will typically provide 300 of Spot Check measurements at an ambient temperature of 25°C.

Li-Poly battery, Model No. CT-2500:
Li-Poly battery, 3.7V, 2500mAh, charging time 5-6 hours and will typically provide 300 of Spot Check measurements at an ambient temperature of 25°C.

Use ONLY AC power supply, Model No. FW 7660M/06
Medical power supply with option for country-specific input plug
Input: 100-240V AC / 50-60Hz / 250mA
Output: 6V DC / 1.4 A

Storing Data:
Data memory on device:
12 000 Spot Check measurements can be stored in storage.

Communication interface:
USB 2.0
ONLY connect devices via USB which provide safety extra low voltage 5V DC and double isolation according to IEC 60601-1 or IEC 60950-1

VM-2400® - Counter Key:
Only connect the Viamed VM-2400® - Counter Key.

Miscellaneous:
- Spot Check measurement mode
- Construction: class IP31 degree of protection against foreign bodies and water provided the device is in the intended position of use with the display facing the top while placed on the back
- Classification: Class IIa device, in accordance with MDD 93/42/EEC
- Electrical safety: Class of protection II / Type BF (VM-2400-Sensor)
- Dimensions (L x W x H): 15,5 x 7,5 x 3,5 cm
- Weight: approx. 550 g (with batteries, without sensor)

Note: This product complies with ISO 10993-1, Biological Evaluation of Medical Devices
Part 1: Evaluation and Testing.

Note: This device is not made with natural rubber latex.

Applied Standards:
IEC 60601-1 (3rd Ed); IEC 60601-1 (2nd Ed); IEC 60601-1-2 (4th Ed);
EN ISO 80601-2-61; EN ISO 9919
ISO 14971

# Testing summary

SpO2 Accuracy:
To validate the SpO2 accuracy controlled hypoxia studies were conducted on a pool of consenting subject volunteers at an independent research laboratory. A pool of 12 healthy, female and male with skin tones ranging from light to dark subjects that are 18 years of age and older were studied over two days, with 20-25 arterial blood samples each. All subjects were in good health at the time of the study. The values for COHb and MetHb for all subjects were low. The COHb values were in the range of 0.3 % to 1.2 % and the MetHb are in the range of 0.1 % to 0.6 %. The values for ctHb for all subjects were over 10 g/dl.
In the range of 60 – 100% the measured SpO2 value (arterial blood saturation) of the sensor is compared to the reference SaO2-value of the blood gas analyser samples.
The accuracy is determined and stated in terms of the root-mean-square (Arms) difference between measured values (SpO2) and reference values (SaO2) for all subjects and relates to the ISO 8060-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment.

Hb Accuracy:
To validate the Hb accuracy controlled haemoglobin studies were conducted on a pool of consenting subject volunteers at an independent blood donor service. The pool consists of healthy, female and male that are 18 years of age and older. In the range of 8 – 20 g/dl more than 390 measured Hb value (total haemoglobin) of the sensor were compared to the reference cHb-value of the blood gas analyser samples.
The accuracy is determined and stated in terms of the root-mean-square (Arms) difference between measured values (Hb) and reference values (Hb) for all subjects and relates to the ISO 8060-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment.

Measurement Results:
- Arms: 1.09[7]
  [7] Arms represents appr. 68% of measurements.
- Bias: -0.17 g/dl
- Precision: 0.82 g/dl

MetHb Accuracy:
To validate the MetHb accuracy controlled methemoglobin studies were conducted on a pool of consenting subject volunteers at an independent research laboratory. The pool consists of 6 healthy, female and male with skin tones ranging from light to dark subjects that are 18 years of age and older. In the range of 0 – 15 % the measured MetHb value (methemoglobin concentration) of the sensors is compared with the reference MetHb value (methemoglobin concentration) of blood gas analyser samples.
The determination of accuracy and is be stated in terms of the root-mean-square (Arms) difference between measured values (MetHb) and reference values (MetHb) for all subjects and relates to the ISO 80601-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment

# Packing List – Accessories and Replacement Parts

Packing List:
- VM-2400, main unit
- VM-2400 - Counter Key (500 units)
- VM-2400 - Sensor
- USB data cable
- Power supply (FW7660M/06)
- Power supply adapter (Europe plug)
- Li-Poly battery (CT-2500)
- 4 x Batteries (AA Mignon)
- Silicon protective cover
- Anti-stress ball
- User instruction
- Short instruction
- Hard case carry case

Order Number:
0012400

Accessories and Replacement Parts

VM-2400 - Sensor, REF 0012405

VM-2400 - Counter Key, REF 0012406, 500 available measurements

VM-2400 - Counter Key, REF 0012407, 2500 available measurements

VM-2400 - Counter Key, REF 0012408, 5000 available measurements

Anti-stress ball, REF 0022401

Philips InfraCare® IR-A radiator, REF 0022402

Holder complete, REF tbd

Universal Mounting Kit, REF: 4420601 V-adapter with female pole-mount thread

Universal Mounting Clamp, REF: 0121200, Adapter with vertical and horizontal adjustment

Functional protective case, REF: tbd, Open access to all ports and buttons - with shoulder strap and stand function

Soft carry case, REF 0022170, Carrying bag for main unit and sensor, with shoulder strap “yellow”

Transportation case, REF tbd, Hard cover carrying case for main unit, sensors and disposables

VM-2400 Silicone Protective Cover, REF tbd

Power supply, FW7660M/06, REF tbd, power supply for continuous operation of the VM-2400 and charging of Li-Poly battery, Input 100-240V AC / 50-60Hz, Output: 6V DC / 1.4A

Europe power supply plug, REF 4430604, Europe adapter for the power supply FW7660M/06

UK power supply plug, REF 4430603, UK adapter for the power supply FW7660M/06

AUS power supply plug, REF tbd, Australian adapter for power supply FW7660M/06

USA/JP power supply plug, REF tbd, USA/Japan adapter for power supply FW7660M/06

Li-Poly battery, CT-2500, REF 4420590, special rechargeable battery for use with the VM-2400, 3.7V / 2500mAh