QC 22 Non Conformance Report

Date	08/09/2016
Issue id	78700
BSI Ref (if applicable)	1368854N3 / Manufacture and test: Nerve stimulators
Responsibility	Derek Lamb
Non-Conformance	“Supplier returns are not automatically handled through the formal CAPA/actions process. Although the client often requests investigations and corrective action from suppliers this it not managed formally through the CAPA process. “
Investigation By:	Derek Lamb
Investigation Issue id / (if applicable)
Corrective Action By:	Derek Lamb
Corrective Action Issue ID (if applicable):	N/a
Time Scale for Corrective Action	Done
Corrective Action:	78702. / System has been updated to automate the trigger for all returns so a decision is logged as to Raising a QC21 or QC22 Form. / Automated issue to point to location failure information can be retrived by referencing the internal returns BOX number. Also provides reference to VM3COP10 Customer complaints and Vigilance procedures. / See Issue#78621 / For example run through
Follow-up future issue id / (Effectiveness verification)	The rolling supplier return audit (69), is triggered every 2 months / the following notes have been added / “ / Check to see if any returns have shipped since the 08th September 2016, / Check to see if a non conformance was raised against the BOX number, search Issues history for the BOXID / If found can close non conformance issue ID78700, and then remove these notes from the audit. / if non found leave these notes until the next Audit” / This is next due to get actioned/created on the 26 Oct 2016.
Effectiveness verification
Closed By:
Closed on