VersaStream Sampling Line Test

Test Protocol No.: TR-11-01-0001-17

| Test location: | bluepoint MEDICAL GmbH & Co. KG An der Trave 15 D – 23923 Selmsdorf, Germany |
| --- | --- |
| Test date: | 2017-01-04 |
| Report date: | 2017-01-05 |

Responsible for the technical content of the test report:

Examiner:			R. Jürgensen

______________________________

COMPANY CONFIDENTIAL

This document is the property of Bluepoint MEDICAL GmbH & Co. KG and is protected by German and international copyright law, and may not, without the written permission by Bluepoint MEDICAL GmbH & Co. KG, be given to any third parties.

This report may only be reproduced or published in its entirety. Extracts could adulterate the stipulations and the results.

Table of Content

1. Summary	3

2. Equipment under Test (EuT)	4

3. Method	5

3.1 Measurement Accuracy	5

3.2 Negative pressure Test	8

4. Results	10

4.1 Measurement Accuracy	10

4.2 Negative pressure Test	16

5. Conclusion	17

5.1 Measurement Accuracy	17

5.2 Negative pressure Test	17

# 1. Summary

| Purpose: | The purpose of this test is to analyze and verify the accuracy and functional compatibility of VersaStream sampling lines to the original Nomoline sampling lines. |
| --- | --- |
| Method: | Measure and verify the CO2 and RR (Respiratory Rate) measurement accuracy of the VersaStream sampling lines with a CO2 respiration simulator. The CO2 measurement accuracy specification: RMS < 0,2. The RR measurement accuracy specification: RMS < 2. Measure and verify the negative pressure of the VersaStream sampling lines with a precision digital manometer. The negative pressure accuracy specification: < 25 mbar |
| Results: | The VersaStream sampling line test is passed. |
| Observations: | The CO2 measurement accuracy is: RMS < 0,2. The RR measurement accuracy is: RMS < 2. The negative pressure accuracy are: < 25 mbar |
| Conclusion: | The CO2 and RR test verifies that the accuracy of the tested VersaStream sampling lines is observed. The Pressure Test shows no significant differences between the Nomoline and the VersaStream sampling lines. |

# 2. Equipment under Test (EuT)

| Manufacturer | Name | Lot Number |
| --- | --- | --- |
| Masimo Sweden AB | NomoLine | 15020096 |
| Bluepoint medical GmbH & Co. KG | VersaStream Short Term | 2618.12.16 |
| Bluepoint medical GmbH & Co. KG | VersaStream Long Term | 2711.12.16 |

Photos of EuT

Nomoline with Airway Adapter

VersaStream Short Term with Airway Adapter

VersaStream Long Term with Airway Adapter

# 3. Method

## 3.1 Measurement Accuracy

3.1.1 Test equipment

CapnoCheck , CO2-Respiratory-Simulator		Test gas 3; 5; 7,5 Vol% + 0,1 Vol%

Test

object

IRMA

UART-USB-Converter for mainstream CO2	UART-USB-Converter with integrated IRMA-Sensor as reference				side stream CO2 ISA-module

3.1.2 Test Setup

Test set-up

Software Gas Master V3.0 user interface

Test description

Use the Gas Master V3.0 software to record the measurement data.

12 respiratory rate steps are simulated with 3 CO2 test gas concentrations a according to table:

| Parameter | Parameter |
| --- | --- |
| Co2 test gas | 3Vol%, 5Vol%, 7,5Vol%, |
| Respiratory rate | 10, 20, 30, 40, 50, 60, 70, 80, 90, 100, 110, 120 |

Connect the UART-USB mainstream converter with a PC and the IRMA CO2 sensor for reference CO2 data. Connect the UART-USB side stream converter with integrated ISA-CO2 module with a PC and the EuT.

Connect the IRMA and the EuT with the outlet valve of the CapnoCheck. Start the CapnoCheck with the parameter in the table above.

Each respiratory rate test step has duration 120 seconds.

In the first 60 seconds the values and simulator are allowed to stabilize. Thereafter the respiratory rate and CO2 value are recorded

The functional performance criteria required to pass the test:

- CO2  measurement accuracy: 	RMS*  < 0,2

- RR measurement accuracy: 		RMS  < 2

*RMS: Root Mean Squared difference

RMS=i=1n(UUT measured i-Target i)2n

UUT: Measured value

Target: Target value

The recorded raw data were additional illustrated as a capnogram.

Ambient temperature:		22°C

Rel. Humidity:		34%

## 3.2 Negative pressure Test

3.2.1 Test equipment

Test

object

Digital Manometer with adapter		UART-USB-Converter for side stream

CO2 ISA-Sensor

3.2.2 Test Setup

Test set-up

3.2.3 Test description

Connect the UART-USB side stream converter with integrated ISA-CO2 module with a power supply. Switch on the digital manometer.

In the first 60 seconds the values are allowed to stabilize. Thereafter the momentary pressure value is recorded.

The functional performance criteria required to pass the test:

- Negative pressure accuracy: 	< 25 mbar

Ambient temperature:		22°C

Rel. Humidity:		34%

# 4. Results

## 4.1 Measurement Accuracy

4.1.1 CO2 Accuracy

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 3 Vol%

|  |  |  | VersaStream | VersaStream |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CapnoCheck | IRMA | NomoLine | Short Term |  | Long Term |  |
| RR [bpm] | et CO2 | et CO2 | et CO2 |  | et CO2 |  |
| 10 | 3,0 | 3,0 | 2,9 |  | 3,0 |  |
| 20 | 3,0 | 3,0 | 2,9 |  | 3,0 |  |
| 30 | 3,0 | 3,0 | 2,9 |  | 3,0 |  |
| 40 | 3,0 | 3,0 | 2,9 |  | 3,0 |  |
| 50 | 3,0 | 3,0 | 2,9 |  | 3,0 |  |
| 60 | 3,0 | 3,0 | 2,8 |  | 3,0 |  |
| 70 | 3,0 | 3,0 | 2,8 |  | 3,0 |  |
| 80 | 3,0 | 3,0 | 2,8 |  | 3,0 |  |
| 90 | 3,0 | 3,0 | 2,8 |  | 2,9 |  |
| 100 | 3,0 | 2,9 | 2,8 |  | 2,9 |  |
| 110 | 3,0 | 2,9 | 2,8 |  | 2,9 |  |
| 120 | 2,9 | 2,9 | 2,8 |  | 2,8 |  |
| RMS | 0,029 | 0,050 | 0,165 |  | 0,076 |  |

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 5 Vol%

|  |  |  | VersaStream | VersaStream |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CapnoCheck | IRMA | NomoLine | Short Term |  | Long Term |  |
| RR [bpm] | et CO2 | et CO2 | et CO2 |  | et CO2 |  |
| 10 | 5,0 | 5,0 | 5,0 |  | 5,0 |  |
| 20 | 5,0 | 5,0 | 4,9 |  | 5,0 |  |
| 30 | 5,0 | 5,0 | 4,9 |  | 4,9 |  |
| 40 | 5,0 | 5,0 | 4,9 |  | 5,0 |  |
| 50 | 5,0 | 5,0 | 4,9 |  | 4,9 |  |
| 60 | 5,0 | 5,0 | 4,8 |  | 5,0 |  |
| 70 | 5,0 | 4,9 | 4,9 |  | 5,0 |  |
| 80 | 5,0 | 4,9 | 4,8 |  | 4,9 |  |
| 90 | 5,0 | 4,9 | 4,8 |  | 4,9 |  |
| 100 | 4,9 | 4,8 | 4,8 |  | 4,9 |  |
| 110 | 4,9 | 4,8 | 4,8 |  | 4,9 |  |
| 120 | 4,9 | 4,8 | 4,7 |  | 4,8 |  |
| RMS | 0,050 | 0,112 | 0,168 |  | 0,091 |  |

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 7,5 Vol%

|  |  |  | VersaStream | VersaStream |  |
| --- | --- | --- | --- | --- | --- |
| CapnoCheck | IRMA | NomoLine | Short Term | Long Term |
| RR [bpm] | et CO2 | et CO2 | et CO2 | et CO2 |
| 10 | 7,5 | 7,5 | 7,5 | 7,4 |
| 20 | 7,5 | 7,5 | 7,4 | 7,4 |
| 30 | 7,5 | 7,5 | 7,4 | 7,4 |
| 40 | 7,5 | 7,4 | 7,4 | 7,4 |
| 50 | 7,5 | 7,4 | 7,4 | 7,4 |
| 60 | 7,5 | 7,4 | 7,3 | 7,4 |
| 70 | 7,5 | 7,4 | 7,3 | 7,3 |
| 80 | 7,5 | 7,4 | 7,3 | 7,3 |
| 90 | 7,4 | 7,3 | 7,3 | 7,3 |
| 100 | 7,4 | 7,3 | 7,2 | 7,3 |
| 110 | 7,4 | 7,3 | 7,2 | 7,2 |
| 120 | 7,4 | 7,2 | 7,2 | 7,2 |
| RMS | 0,058 | 0,147 | 0,197 | 0,180 |

Capnogram

Respiratory Rate (RR): 10 bpm

CO2 test gas concentration: 5 Vol%

Capnogram

Respiratory Rate (RR): 50 bpm

CO2 test gas concentration: 5 Vol%

Capnogram

Respiratory Rate (RR): 100 bpm

CO2 test gas concentration: 5 Vol%

4.1.2 Respiratory Rate (RR) Accuracy

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 3 Vol%

| CapnoCheck | IRMA | NomoLine | VersaStream Short Term | VersaStream Long Term |
| --- | --- | --- | --- | --- |
| RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] |
| 10 | 10 | 10 | 10 | 10 |
| 20 | 19 | 20 | 19 | 20 |
| 30 | 30 | 30 | 30 | 30 |
| 40 | 40 | 40 | 39 | 40 |
| 50 | 50 | 50 | 50 | 50 |
| 60 | 59 | 59 | 60 | 60 |
| 70 | 70 | 70 | 69 | 69 |
| 80 | 80 | 80 | 80 | 80 |
| 90 | 90 | 90 | 90 | 90 |
| 100 | 99 | 99 | 100 | 100 |
| 110 | 110 | 110 | 109 | 109 |
| 120 | 120 | 120 | 119 | 119 |
| RMS | 0,50 | 0,41 | 0,64 | 0,50 |

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 5 Vol%

| CapnoCheck | IRMA | NomoLine | VersaStream Short Term | VersaStream Long Term |
| --- | --- | --- | --- | --- |
| RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] |
| 10 | 9 | 10 | 10 | 9 |
| 20 | 20 | 20 | 19 | 20 |
| 30 | 30 | 30 | 30 | 30 |
| 40 | 40 | 40 | 40 | 40 |
| 50 | 49 | 50 | 50 | 50 |
| 60 | 60 | 60 | 59 | 60 |
| 70 | 70 | 70 | 70 | 70 |
| 80 | 80 | 80 | 80 | 80 |
| 90 | 90 | 89 | 89 | 89 |
| 100 | 101 | 100 | 100 | 100 |
| 110 | 110 | 110 | 109 | 110 |
| 120 | 121 | 119 | 119 | 120 |
| RMS | 0,58 | 0,41 | 0,64 | 0,41 |

Respiratory Rate (RR): 10 bpm till 120 bpm

CO2 test gas concentration: 7,5 Vol%

| CapnoCheck | IRMA | NomoLine | VersaStream Short Term | VersaStream Long Term |
| --- | --- | --- | --- | --- |
| RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] | RR [bpm] |
| 10 | 10 | 10 | 10 | 10 |
| 20 | 20 | 19 | 20 | 19 |
| 30 | 30 | 30 | 29 | 30 |
| 40 | 40 | 40 | 39 | 40 |
| 50 | 50 | 49 | 49 | 49 |
| 60 | 60 | 60 | 59 | 60 |
| 70 | 69 | 70 | 70 | 70 |
| 80 | 80 | 80 | 80 | 80 |
| 90 | 90 | 89 | 89 | 89 |
| 100 | 100 | 100 | 100 | 100 |
| 110 | 110 | 109 | 109 | 109 |
| 120 | 119 | 119 | 119 | 120 |
| RMS | 0,41 | 0,64 | 0,76 | 0,57 |

## 4.2 Negative pressure Test

| Sample volume | NomoLine | VersaStream Short Term | VersaStream Long Term |
| --- | --- | --- | --- |
| 60ml/min | 19 mbar | 20 mbar | 23 mbar |

Diagram

# 5. Conclusion

## 5.1 Measurement Accuracy

5.1.1 CO2 Accuracy

The CO2 Test verifies that the accuracy of the tested VersaStream sampling lines is observed. The Root Mean Square is < 0,2.

5.1.2 Respiratory Rate (RR) Accuracy

The Respiratory Rate Test verifies that the accuracy of the tested VersaStream sampling lines is observed. The Root Mean Square is < 2.

## 5.2 Negative pressure Test

The Pressure Test shows no significant differences between the Nomoline and the VersaStream sampling lines. The pressure difference is < 25 mbar.