| Dimensions and housing | To suit Stealth CDLSE housing |
| --- | --- |
| Measuring Range | 0-2 ATM PO2 (with over-range capability to 2.3 ATM PO2) |
| Accuracy at constant temperature, pressure and humidity. | Accuracy – Within + / - 1% of reading (or ±2 mBar Oxygen whichever is greater) up to specified range, when using temperature compensation circuit (or a fixed 100 Ohm resistor at 25°C) and calibrated using 100% oxygen at known atmospheric pressure and 25°C Acceptable limits are Error mBar pO2 mBar pO2 Min Max 209 -4 4 1000 0 0 Calibration point 1500 -15 15 2000 -40 40 Specified Range 2300 -100 100 Over-range capability |
| Offset Voltage | Less than 0.5% oxygen equivalent at 25 °C in zero gas after 36 seconds |
| Temperature Performance | Better than 0.05% oxygen /°C over operating temperature range (Ie when calibrated at 20°C, an increase to 40°C or a decrease to 0°C could result in an error of 1.0% oxygen). |
| Pressure Effects | Sensor is sensitive to partial pressure of oxygen, and therefore its output is linearly affected by atmospheric pressure |
| Operating Temperature | 0 to 40°C |
| Storage Temperature | -30°C to 70°C |
| Humidity | 0-99% RH non condensing |
| Magnetic Effect | Maximum magnetic field of 5 nano Tesla at 10 centimetres (NATO STANAG 2897 Class A) |
| Response Time | Previously agreed 6s T90. Will accept 10s for T90 if this improves linearity to 2 Bar PO2. |
| Lifetime | 2 to 3 years in air at 25°C |
| Connector | Bare Wires |
| Electrical Output | 60-100 Microamperes in air @ 25 °C sea level |
| Required Load | 10K tuned by thermistor compensation network (nominal 120R) to generate 7-13mV in air at 25°C (sea level), |
| Packaging | To be Supplied in air tight containers or bags |
| Shelf Life in original packaging | Better than 6 months, followed by normal 3 year life as defined above |
| Weight | To be confirmed |
| Shock | To withstand a drop from 0.1m onto a hard wooden surface, and for the cell output to recover back to its original reading (+/-2% of reading) within 10 seconds |
| Orientation effects | Less than 0.2% Oxygen in any orientation |
| Labelling | To be Confirmed |