## Sheet1
Date | Barometric Pressure | W/O | Assy serial # | Cable serial # | Cell Serial # | Cell type | Temp (K) | mV Reading in Air | PPO2 Reading @ 100% O2 | mV reading in 100% O2 | PPO2 Reading @ 100% O2 - after calibration | mV Reading @ 100% O2 - after calibration | PPO2 reading in air | mV reading in air | mV Reading @ 100% O2 | PO2 Reading @ 100% O2 | 0.5 BarG mV Reading @ 100% O2 | 0.5 BarG PO2 Reading @ 100% O2 (1.40 - 1.60) | 1 BarG mV Reading @ 100% O2 | 1 BarG PO2 Reading @ 100% O2 (1.90 - 2.10) | 1.3 BarG mV Reading @ 100% O2 | 1.3 BarG PO2 Reading @ 100% O2 (2.20 - 2.40) | 1.5 BarG mV Reading @ 100% O2 | 1.5 BarG PO2 Reading @ 100% O2 (2.40 - 2.60) | mV Reading @100% O2 | PO2 Reading @100% O2 | mV reading in air | PPO2 reading in air
2017-04-19 | 1.018 | DEV | Dev 4 |  | 1 | PROTO | 292.8 |  |  |  | 1.018 | 25.4 | 0.215 | 5.39 | 25.38 | 1.018 | 32.62 | 1.508 | 49.59 | 1.988 | 56.76 | 2.276 | 61.62 | 2.47 | 25.43 | 1.019 | 5.42 | 0.216
[5 empty] | 2 | [5 empty] | 1.018 | 25.4 | 0.215 | 5.4 | 25.4 | 1.019 | 32.53 | 1.508 | 49.32 | 1.977 | 56.38 | 2.26 | 61.22 | 2.455 | 25.53 | 1.023 | 5.43 | 0.217
[5 empty] | 3 | [5 empty] | 1.018 | 25.39 | 0.215 | 5.37 | 25.4 | 1.018 | 32.6 | 1.508 | 49.53 | 1.986 | 56.68 | 2.273 | 61.55 | 2.468 | 25.43 | 1.019 | 5.41 | 0.217
|  | Test Stage | [8 empty] | 1 |  | 2 |  |  | 3 |  | 4 |  | 5 |  | 6 |  | 7 |  | 8 |  | 9
| 1.018 | Target Value | [8 empty] | 1.018 |  | 0.21378 |  |  | 1.018 |  | 1.518 |  | 2.018 |  | 2.318 |  | 2.518 |  | 1.018 |  | 0.21378
|  | Max Allowable Value | [8 empty] | 1.0282 |  | 0.21778 |  |  | 1.0282 |  | 1.533 |  | 2.058 |  | 2.418 |  | 2.618 |  | 1.0282 |  | 0.21778
|  | Min Allowable Value | [8 empty] | 1.0078 |  | 0.20978 |  |  | 1.0078 |  | 1.503 |  | 1.978 |  | 2.218 |  | 2.418 |  | 1.0078 |  | 0.20978
|  | Allowable Error | [8 empty] | 0.01018 |  | 0.004 |  |  | 0.01018 |  | 0.015 |  | 0.04 |  | 0.1 |  | 0.1 |  | 0.01018 |  | 0.004
2017-04-20 | 1.0145 | DEV | Dev 4 |  | 4 | PROTO | 292 |  |  |  | 1.015 | 5.31 | 0.215 | 5.37 | 25.36 | 1.017 | 37.43 | 1.5 | 49.24 | 1.973 | 56.19 | 2.253 | 60.76 | 2.435 | 25.47 | 1.02 | 5.43 | 0.217
[5 empty] | 5 | [5 empty] | 1.015 | 5.31 | 0.215 | 5.38 | 25.32 | 1.015 | 37.34 | 1.498 | 49.13 | 1.97 | 56.16 | 2.252 | 60.8 | 2.438 | 25.48 | 1.021 | 5.42 | 0.217
[5 empty] | 6 | [5 empty] | 1.015 | 5.31 | 0.214 | 5.32 | 25.3 | 1.015 | 37.46 | 1.502 | 49.41 | 1.981 | 56.44 | 2.263 | 61 | 2.447 | 25.32 | 1.015 | 5.33 | 0.214
| 1.0145 | Target Value | [8 empty] | 1.0145 |  | 0.21304 |  |  | 1.0145 |  | 1.5145 |  | 2.0145 |  | 2.3145 |  | 2.5145 |  | 1.0145 |  | 0.21304
|  | Max Allowable Value | [8 empty] | 1.0246 |  | 0.21704 |  |  | 1.0246 |  | 1.5295 |  | 2.0545 |  | 2.4145 |  | 2.6145 |  | 1.0246 |  | 0.21704
|  | Min Allowable Value | [8 empty] | 1.0044 |  | 0.20904 |  |  | 1.0044 |  | 1.4995 |  | 1.9745 |  | 2.2145 |  | 2.4145 |  | 1.0044 |  | 0.20904
|  | Allowable Error | [8 empty] | 0.010145 |  | 0.004 |  |  | 0.010145 |  | 0.015 |  | 0.04 |  | 0.1 |  | 0.1 |  | 0.010145 |  | 0.004