| Measurement Range: | 0 % ... 100 % oxygen |
| --- | --- |
| Output in Ambient Air: | 6.5 to 9 mV |
| Electrical Interface: | Flying leads with 3 pin female connector (Molex® 22-01-2037) |
| Repeatability: | < 1% vol. O2 @ constant temperature and pressure |
| Linearity error: | 0.1 to 1.0 bar ppO2 < 3% relative |
|  | 0.1 to 2.5 bar ppO2 < 5% relative (at ppO2 > 1.0 bar after 1 h equilibration time) |
| Accuracy at constant temperature, pressure and humidity | Accuracy – Within + / - 3% of reading (or ±5 mBar Oxygen whichever is greater) up to specified range, when calibrated using 100% oxygen at known atmospheric pressure and 25°C Acceptable limits are Error mBar pO2 mBar pO2 Min Max 209 -4 4 1000 0 0 Calibration point 1500 -30 30 2000 -75 75 Specified Range 2300 -100 100 Over-range capability Hysteresis after use at high pressure according to accuracy specification (Within + / - 3% of reading) |
| Response time: | < 5 sec. to 90% of final value |
| Zero Offset Voltage: | < 0.5% oxygen reading in 100% nitrogen @ 25°C after 36 seconds |
| Cross Interference: | < 0.5% vol. O2 response to: 5% CO2 balance N2 |
| Influence of Humidity: | - 0.03% rel. per % RH at 25°C |
| Pressure Range: | 600 to 2500 mbar |
| Influence of Mechanical Shock: | < 1% relative after a fall from 1m |
| Operating Temperature: | 0°C to 50°C |
| Temperature Compensation: | Built-in NTC compensation |
| Effect of Temperature Compensation (steady state): | Between 0°C and +50°C: 5% relative error |
| Operating Humidity: | 0 - 99% RH non-condensing |
| Long Term Output Drift: | < 1% vol. oxygen per month, Typically < -15% relative over lifetime |
| Storage Temperature: | -20°C to +50°C |
| Recommended Storage: | +5°C to +15°C |
| Recommended Load: | ≥10 kOhms |
| Warm-Up Time: | < 30 minutes, after replacement of sensor |
| Nominal Sensor Lifetime: | 500,000% vol oxygen hours |
| Weight: | Approximately 28 grams |
| Warranty Period: | 15 months (including 3-month shelf life) |
| Typical Lifetime | 20 to 24 months from shipment (including shelf life) in air at 25°C. |
| Electrical output | 30-60 Microamperes in air @ 25 °C sea level |
| Packaging | In plastic bag |
| Part No.: | OOD103-JFD: 1002461 (REF DB200381) |
| All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air. |  |

Mechanical draft:

General tolerances ISO 2768-c

General tolerances ISO 2768-c

Customized Design Specification:

- Integration of additional Zitex membrane in sensor gas diffusion way

- using of non magnetic components for the sensor

- Conformal coating on the sensor PCB with PLASTIK 70

- 2 x 2mm holes drilled in the sensor PCB

- No information about EnviteC on PCB top / bottom sides, information for ordering process and traceability only.

PCB Layout:

| Top | Bottom |
| --- | --- |

Product Labeling

| OOD103-JFD: 1002461 (REF DB200381) |  |
| --- | --- |
| Label on Sensor (90x15) | Label on Blister- Card (63,5x46,6) |

Customer Approval

Date:

Signature: