[image: image3.jpg]




Specification Datasheet - 8010050

JFD Ltd.
DB200381 Oxygen Sensor
[image: image4.png]2276+0.2

T Molex 22-01-2037

)

not used
; ; I )

50%2

@27.2+0.2

204+0.2*





v1.7
18th July 2018
	Specifications


	VST part number
	8010050

	OEM model number
	DB200381 Oxygen Sensor

	Output in ambient air
	6.5 – 9.0 mV

	Electrical output
	30 to 60 µA

	Output connector
	Flying leads with 3 pin female connector - Molex 22-01-2037

	Input
	Open input – see drawing and images

	Input O ring
	Not applicable

	Range
	0 – 100% oxygen

	Pressure range
	600 to 2500 mbar

	Repeatability
	<1% volume oxygen at constant temperature and pressure

	Linearity error
	0.1 to 1.0 bar < 3% relative

0.1 to 2.5 bar < 5% relative

(at ppO2 > 1.0 bar after 1-hour equilibration time)

	Accuracy at constant temperature, pressure and humidity

	Accuracy – Within  + / - 3% of reading  (or ±5 mBar oxygen, whichever is greater) up to specified range,  when calibrated using 100% oxygen at known atmospheric pressure and 25°C

Acceptable limits are 



Error mBar pO2



mBar pO2
  Min
Max


  209

    -4
   4


1000

     0
   0
Calibration point 

1500

  -30
  30


2000

  -75
  75
Specified range

2300

-100
100
Over-range capability

Hysteresis after use at high pressure according to accuracy specification  (within  + / - 3% of reading)  

	Warm up time
	< 30 minutes after sensor installed 

	Zero offset voltage
	< 0.5% oxygen reading in 100% nitrogen @ 25ºC after 36 seconds

	Cross interference
	<0.5% volume oxygen response to 5% CO2 balance N2

	Response time
	< 5 s for 90% of final value.

	Operating humidity
	0 – 99% R.H. (non-condensing).

	Influence of humidity
	- 0.03% relative per % R.H. at 25 ºC.

	Influence of mechanical shock
	< 1% after fall from 1m

	Operating temperature range
	0 – 50 ºC

	Storage temperature
	-20 to +50 ºC

	Recommended storage temperature
	+5 to +15 ºC

	Temperature compensation
	Integrated NTC network

	Temperature compensation error (steady state)
	0 °C to 50 °C, 5% relative error

	Nominal sensor life
	500,000 % oxygen hours

	Typical sensor life
	20 to 24 months from date of invoice (including shelf life).

	Long term output drift
	< 1% volume oxygen per month, typically < -15% over sensor lifetime.

	Load required
	≥10 K Ohms


Page 2 of 4

	Input - membrane protection grille
	No

	Input - hydrophobic membrane
	Yes

	Conformal coating on PCB
	Yes

	Pressure equalization - PCB
	Yes 

	Pressure equalization - output cap
	Not applicable

	Weight (approximate)
	28 g

	Flow diverter
	No

	Packaging
	Gas barrier bag

	Warranty period
	12 months from date of sales invoice

	Application specific components
	Low/non-magnetic components utilized: Cathode, internal wiring. Suitability and QA assessment to be carried out by JFD prior to sensor production.

	All specifications are applicable at standard conditions: 1013 mbar, 25°C dry ambient air


	[image: image5.png]Vandagraph
>ensor Technologies




Drawings/images


 Dimensions in mm
*Sensor housing height, without PCB
Page 3 of 4
	Labelling


Oxygen sensor label 90.0mm x 15.0mm

White vinyl label, text & graphics - black

[image: image1.jpg]JFD Ltd.
Enterprise Drive
Westhih,
Aberdeen
AB326TQ
United Kingdom

Oxygen Sensor

A

X

PIN: DB200381

il
2017-07

S/N: 123456




Gas barrier bag label 63.5mm x 46.6mm

White label, text & graphics – black

[image: image2.jpg]JFD Ltd.
\ Enterprise Drive

Westhil,

Averdeen

AB326TQ
United Kingdom

Oxygen Sensor  P/N: DB200381

M 2017-07 1] x

SIN 123456

R e

This oxygen sensor contains caustic liquid, in case of leakage avoid contact with eyes
and skin. Do not puncture or staple this bag. Do not remove the oxygen sensor from
this bag until ready to use. Depending upon the ambient conditions the oxygen sensor
may require 30 minutes to reach signal stability after installation.





Page 4 of 4

Vandagraph Sensor Technologies Ltd.

15 Station Road, Cross Hills, Keighley, West Yorkshire BD20 7DT, United Kingdom

Tel +44 (0) 1535 634900   Fax +44 (0) 1535 635582   e-mail: info@vandagraphst.com

www.vandagraphst.com

