| DB200381 Oxygen Sensor Specification Document Number: DB200-SP-5223 Revision: 01 |  |
| --- | --- |

ISSUED FOR CONSTRUCTION 03-10-2018.

PO: 11807111

| Document Information |  |  |  |  |
| --- | --- | --- | --- | --- |
| Number | Title | Classification | Current Revision | Date |
| DB200-SP-5223 | DB200381 Oxygen Sensor Specification | Commercial In Confidence | 01 | 08/08/18 |

| Revision History |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Rev | Date | BY | CHK | APP | Comments |
| 0 | 2nd April 2018 | MH | SB | SC | Initial Issue |
| 1 | 8th August 2018 | MH | JM | SB | ECN: 22688 Updated dimensions |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |

| Original Approvals |  |  |  |
| --- | --- | --- | --- |
| Responsibility | Name | Position | Date |
| Author | Michael Hossack | Senior Electronics Engineer | 02/04/18 |
| Checked | Stephen Blatherwick | Senior Development Engineer | 02/04/18 |
| Approval | Steven Coull | Product Engineering Manager | 02/04/18 |

Copyright Details

© 2018 JFD Copyright of this document is the property of JFD and it may not be copied, used or otherwise disclosed in whole or in part except with prior written permission from JFD or, if this document has been furnished under a contract with another party, as expressly authorised under that contract.

# Introduction

This document describes the specification for the DB200381 non-magnetic oxygen sensor intended to be installed within JFD’s range of electronic rebreathers.

# Specification

| Output in ambient air | 6.5 to 9.0 mV |
| --- | --- |
| Electrical output | 30 to 60 µA |
| Output connector | Flying leads or solder pads |
| Input | Open input Hydrophobic membrane |
| Input O ring | Not applicable |
| Output range | 0 to 2500 mbar PPO2 |
| Pressure range | 600 to 10000 mbar |
| Repeatability | < 1% volume oxygen at constant temperature and pressure |
| Linearity error | - to 0.4 bar: < 30 mbar > 0.4 to 2.0 bar: < 60 mbar > 2.0 to 2.3 bar: < 100 mbar |
| Warm up time | < 30 minutes after sensor installed |
| Zero offset voltage | < 0.5% oxygen reading in 100% nitrogen @ 25 °C after 36 seconds |
| Cross interference | < 0.5% volume oxygen response to 5% CO2, balance N2 |
| Response time | < 10 seconds for 90% final value |
| Operating humidity | 0 to 99% RH (non-condensing) |
| Influence of humidity | -0.03% relative per % R.H. at 25 °C |
| Influence of mechanical shock | < 1% after fall from 1 m |
| Operating temperature range | 0 to 50 °C |
| Storage temperature range | -30 to +70 °C |
| Temperature compensation error | 0 to 50 °C, 5 % relative error |
| Shelf life in original packaging | > 6 months |
| Operation life | > 2 years in air |
| Long term output drift | < 1% volume oxygen per month |
| Required load | 10 K Ohms |
| Weight | < 30 g |
| Dimensions | See Section 2.1 below |
| Magnetic signature | < 3.5 nT |
| Orientation | <0.2% oxygen in any orientation |
| Packaging | Gas barrier bag |
| Marking | JFD logo and part number (DB200381). Serial number in format to be agreed between JFD and supplier. Any other markings as required by relevant legislation. |

## Dimensions

To be as shown in Figure 1 below.  Any external circuitry to be flush with back of sensor body, except for solder joints which may protrude.

Figure 1 – Sensor Dimensions

## Supplier Testing

All sensors to be supplied to JFD with results of testing the following parameters in accordance with this specification:

- Output in air

- Zero offset voltage

- T90 response time

- Linearity error (0 – 100% oxygen at ambient pressure)